## Guide
DATA COLLECTION TEMPLATE v2 — Restructured for Clarity
SHEET 1: Electrode Materials → One row per material. Check which vendors offer it. Fill properties.
SHEET 2: Liner Materials → Same logic. One row per liner material.
SHEET 3: Vendor Models → One row per vendor model (Rosemount 8700, Promag W, OPTIFLUX, etc.). Fill KPIs.
SHEET 4: JESA Project Data → ALREADY FILLED! 14 flowmeters from PMP CAP Plant with all specs extracted.
SHEET 5: CAPEX & OPEX → One row per vendor model. Fill pricing and lifecycle costs.
SHEET 6: Drift Detection → One row per drift indicator. Fill thresholds and diagnostics info.
SHEET 7: Sources & Evolution → Reference links and material innovation timeline.

## Vendor Models
Vendor | Model / Series | Type | Equivalent Emerson | Equivalent E+H | Equivalent Krohne | Available Electrodes | Available Liners | Accuracy (% reading) | Repeatability (%) | Turndown Ratio | Min Conductivity (µS/cm) | Flow Velocity Range (m/s) | Min Temp (°C) | Max Temp (°C) | Max Pressure (Standard) | Max Pressure (Option) | Pipe Sizes (DN) | Excitation Type | Communication Protocols | Empty Pipe Detection | IP Rating | Transmitter Type | Power Supply | Internal Diagnostics | Source Document | Notes
Emerson | 8705 | Flanged sensor | — | Promag P 300 | OPTIFLUX 4300 | 316L, C-276, C22, Platinum, Tantalum, Titanium | PEX, PTFE, ETFE, PFA, PFA+, PU, Neoprene, Linatex, Adiprene | ±0.25% , ±0.15% (optional) | ±0.1% | >1000:1 | ≥ 5 µS/cm | 0.01–12 m/s | -40°C | +180°C | Class 150 | Class 300, 600, 900, 1500, 2500 | 15–900 mm | Pulsed DC | HART, FOUNDATION Fieldbus, Modbus RS485 | Yes | IP68 | 8732 (integral/remote) or 8712 (wall mount) | AC: 90–250 VAC / DC: 12–42 VDC | Smart Meter Verification, Process Noise Detection, Electrode Coating Detection, Grounding/Wiring Fault, Coil Fault | PDS 00813-0100-4444 | Meilleure accuracy standard (±0.25% vs ±0.5% E+H), seul modèle à offrir Class 2500, 9 liners vs 3 (E+H) et 6 (Krohne)
Emerson | 8711 | Wafer sensor | — | Promag W 10 / P 10 | OPTIFLUX 1100 | 316L, Ni-Alloy, Platinum, Tantalum, Titanium | PTFE, ETFE, PFA, PU, Neoprene | ±0.25% , ±0.15% (optional) | ±0.1% | >1000:1 | ≥ 5 µS/cm | 0.01–9.1 m/s | -40°C | +180°C | Class 150 | Class 300 | 40–200 mm | Pulsed DC | HART, FOUNDATION Fieldbus, Modbus RS485 | Yes | IP68 | 8732 (integral/remote) or 8712 (wall mount) | AC: 90–250 VAC / DC: 12–42 VDC | Smart Meter Verification, Process Noise Detection, Electrode Coating Detection, Grounding/Wiring Fault, Coil Fault | PDS 00813-0100-4444 | Wafer = plus compact et moins cher que W 10 et 1100, mais DN limité à 200 max
Emerson | 8750W | Water sensor | — | Promag W 400 | OPTIFLUX 2100 | 316L, Ni-Alloy | PTFE, Neoprene | ±0.5% , ±0.25% (optional) | ±0.1% | >1000:1 | ≥ 5 µS/cm | 0.01–12 m/s | -20°C | +80°C | Class 150/300 / PN 40 | — | 15 mm–1200 mm | Pulsed DC | HART or Modbus protocol only | Yes | IP68 | Integral or remote (field mount / wall mount) | AC: 90–250 VAC / DC: 12–42 VDC | Smart Meter Verification, Process Diagnostics | PDS 00813-0300-4750 | Moins cher que W 400 mais seulement 2 liners (PTFE, Neoprene) et 2 electrodes (SS, Ni-Alloy). W 400 offre plus de matériaux
Emerson | 8707 | Slurry sensor | — | Promag P 500 | OPTIFLUX 5100 | 316L, C-276, Hastelloy, Platinum, Tantalum, Titanium | PTFE, PFA, PU, Linatex, Neoprene | ±0.25% , ±0.15% (optional) | ±0.1% | >1000:1 | 50 µS/cm when used with 8712H, or 5 µS/cm when used with other transmitters | 0.01–12 m/s | -20°C | +180°C | Class 150 | Class 300 | 80–900 mm | High-Signal Pulsed DC — coils driven at 10 times the current of standard magnetic flow meters | HART, FOUNDATION Fieldbus | Yes | IP68 | 8712H (High-Signal) or 8732 | AC: 90–250 VAC / DC: 12–42 VDC | Smart Meter Verification, Process Noise Detection, Electrode Coating Detection | PDS 00813-0100-4727 | Seul modèle High-Signal (10x coil current) — ni E+H ni Krohne n'ont cette technologie. OPTIFLUX 5100 utilise Ceramic à la place. Obsolète
Emerson | 8750WA | Advanced water | — | Promag W 300 | OPTIFLUX 2300 | 316L, Ni-Alloy C-276, Alloy C22, Platinum, Tantalum, Titanium | PEX, PTFE, ETFE, PFA, PFA+, PU, Neoprene, Linatex, Adiprene | ±0.25% , ±0.15% (optional) | ±0.1% | >1000:1 | ≥ 5 µS/cm | 0.01–12 m/s | -40°C | +180°C | Class 150/300 / PN 40 | — | DN 15–1200 | Pulsed DC | HART, Modbus RS485 | Yes (Tunable EPD) | IP66/IP67/IP68 | 8750WA12 (wall mount) ou 8750WA32 (field mount) | AC: 90–250 VAC / DC: 12–42 VDC | Smart Meter Verification, Ground/Wiring Fault, High Process Noise, Empty Pipe | PDS 00813-0300-4750 (uses 8705/8711/8721 sensors) - PDS 8750WA | Accuracy ±0.25% vs ±0.5% (W 300 et 2300), mais mêmes protocols limités (HART/Modbus) vs full fieldbus chez E+H et Krohne
Emerson | 8721 | Sanitary/hygienic sensor | — | — | — | 316L, Alloy C22, Tantalum, Platinum | PFA | ±0.25% , ±0.15% (optional) | ±0.1% | >1000:1 | ≥ 5 µS/cm | 0.01–12 m/s | -20°C | +150°C | Class 150 | Class 300 | DN 15–200 | Pulsed DC | HART, FOUNDATION Fieldbus, Modbus RS485 | Yes | IP67/IP68 | 8732 (integral/remote) ou 8712 (wall mount) | AC: 90–250 VAC / DC: 12–42 VDC | Smart Meter Verification, Process Noise Detection, Electrode Coating Detection, Grounding/Wiring Fault, Coil Fault | PDS 00813-0100-4444 | Pas d'équivalent direct. E+H a le Promag H (non listé). Seul hygienic chez Emerson
Emerson | 9205 | Next-gen sensor | — | Promag P 500 | OPTIFLUX 4400 | 316L, C-276, C22, Platinum, Tantalum, Titanium | PEX, PTFE, ETFE, PFA, PFA+, PU, Neoprene, Linatex, Adiprene | ±0.25% , ±0.15% (optional) | ±0.1% | >1000:1 | ≥ 5 µS/cm | 0.01–12 m/s | -40°C | +180°C | Class 150 | Class 300 | DN 15–900 | Pulsed DC | HART, Bluetooth | Yes | IP68 | 9732 Magnetic Flow Transmitter with advanced process and device insights, data and device logging capabilities, wireless Bluetooth configuration | AC: 90–250 VAC / DC: 12–42 VDC | 9732 Magnetic Flow Transmitter with advanced process and device insights, data and device logging capabilities, wireless Bluetooth configuration | Emerson product page | Next-gen avec Bluetooth. Pas encore de retour terrain JESA. P 500 et 4400 sont éprouvés
Endress+Hauser | Promag W 10 | Compact water | OPTIFLUX 4400 | — | OPTIFLUX 1100 | 316L, C22, Tantalum | Hard rubber, PU, PTFE | ±0.5% , ±0.2% (optional, Flat Spec) | ±0.1% | >1000:1 | ≥ 5 µS/cm | 0.01 - 10 m/s | -20°C | +90°C | PN 40 / Class 300 | — | DN 25-2400 | Pulsed DC | HART, Modbus RS485 | Yes | IP68 | Compact or remote | DC: 18–30 VDC | Heartbeat Technology | TI01580D | Le plus compact de la série W. Moins de protocols que W 300/400/500 mais suffisant pour applications simples
Endress+Hauser | Promag W 100 | Standard water | 8750W | — | OPTIFLUX 2100 | 316L, Alloy C22, Tantalum | Polyurethane, Hard rubber, PTFE | ±0.5% , ±0.2% (optional, Flat Spec) | ±0.1% | >1000:1 | ≥ 5 µS/cm | 0.01–10 m/s | -20°C | +90°C | PN 40 / Class 300 | — | DN 25–2000 | Pulsed DC | HART, Modbus RS485 | Yes | IP67/IP68 | Compact (integral), Proline 100 | AC: 100–230 VAC / DC: 20–55 VDC | Heartbeat Technology | E+H product page | Entre W 10 et W 400. Web server intégré (W 10 n'a pas). Moins de protocols que W 300
Endress+Hauser | Promag W 300 | Advanced water | 8750WA | — | OPTIFLUX 2300 | 316L, Alloy C22, Tantalum | Polyurethane, Hard rubber, PTFE | ±0.5% , ±0.2% (optional, Flat Spec) | ±0.1% | >1000:1 | ≥ 5 µS/cm | 0.01–10 m/s | -20°C | +90°C | PN 40 / Class 300 | 20K (JIS) | DN 25–2000 | Pulsed DC | HART, PROFIBUS DP/PA, Modbus RS485, EtherNet/IP, PROFINET, Foundation Fieldbus | Yes | IP67/IP68 | Compact or remote, Proline 300, WLAN, web server | AC: 100–230 VAC / DC: 24 VDC | Heartbeat Technology, data logger | E+H product page W300 | Full fieldbus vs 8750WA (HART/Modbus seulement). Même accuracy. Plus de protocols que W 400 (a PROFIBUS PA en plus)
Endress+Hauser | Promag W 400 | Full water | 8750W | — | OPTIFLUX 2100 | 316L, C22, Tantalum | Hard rubber, PU, PTFE | ±0.5% , ±0.2% (optional, Flat Spec) | ±0.1% | >1000:1 | ≥ 5 µS/cm | 0.01 to 10 m/s | -20°C | +90°C | PN 40 / Class 300 | — | DN 25-2000 | Pulsed DC | HART, PROFIBUS DP, Modbus RS485, EtherNet/IP | Yes | IP68 | Compact or remote, WLAN access, web server | AC: 100–230 VAC / DC: 20–55 VDC | Heartbeat Technology, data logger | TI01046D | WLAN access + web server vs 8750W qui n'a ni l'un ni l'autre. Même accuracy ±0.5%
Endress+Hauser | Promag W 500 | Premium water | 8750WA | — | OPTIFLUX 2300 | 316L, C22, Tantalum | Hard rubber, PU, PTFE | ±0.5% , ±0.2% (optional, Flat Spec) | ±0.1% | >1000:1 | ≥ 5 µS/cm | 0.01 to 10 m/s | -20°C | +90°C | PN 40 / Class 300 | — | DN 25-2000 | Pulsed DC | HART, PROFIBUS DP/PA, Modbus RS485, EtherNet/IP, PROFINET, Foundation Fieldbus | Yes | IP68 | Remote only, up to 4 I/Os | AC: 100–230 VAC / DC: 20–55 VDC | Heartbeat Technology, data logger | TI01227D | Remote only, 4 I/Os vs W 400 (compact ou remote, moins de I/Os). Plus de protocols. Plus cher
Endress+Hauser | Promag P 10 | Compact process | 8711 | — | OPTIFLUX 1100 | 316L, C22, Tantalum, Platinum, Titanium | PFA, PFA HT, PTFE | ±0.5% , ±0.2% (optional, Flat Spec) | ±0.1% | >1000:1 | ≥ 5 µS/cm | 0.01 to 10 m/s | -40°C | +180°C | PN 40 / Class 300 | 20K (JIS) | DN 15–600 | Pulsed DC | HART, Modbus RS485 | Yes | IP67 | Compact or remote | DC: 18–30 VDC | Heartbeat Technology | TI01225D | Même matériaux que P 300/P 500 (PFA, Tantalum, Platinum) mais transmitter compact basique. Budget option pour acides
Endress+Hauser | Promag P 100 | Compact process | 8705 | — | OPTIFLUX 4100 | 316L, Alloy C22, Tantalum, Platinum, Titanium, Duplex 1.4462 | PFA, PFA HT, PTFE | ±0.5% , ±0.2% (optional) | ±0.1% | >1000:1 | ≥ 5 µS/cm | 0.01 to 10 m/s | -40°C | +180°C | PN 40 / Class 300 | 20K (JIS) | DN 15–600 | Pulsed DC | HART, Modbus RS485 | Yes | IP67 | Compact (integral), Proline 100, web server | AC: 100–230 VAC / DC: 20–55 VDC | Heartbeat Technology | TI Promag P 100 | Web server + Heartbeat comme P 300 mais sans PROFIBUS/PROFINET. A Duplex 1.4462 comme P 500. Bon compromis
Endress+Hauser | Promag P 200 | Loop-powered process | — | — | — | 316L, Alloy C22, Tantalum, Platinum, Titanium | PFA, PTFE | ±0.5% o.r. ± 2 mm/s | ±0.1% | >1000:1 | ≥ 5 µS/cm | 0.01 to 10 m/s | -25°C | +180°C | PN 40 / Class 300 | — | DN 2.5–300 | Pulsed DC | HART (loop-powered, 2-wire) | Yes | IP67 | Compact, Proline 200, loop-powered (2-wire) | DC: 11.5–42 VDC (loop-powered — c'est sa particularité, il se branche sur la boucle 4-20 mA sans alimentation séparée) | Heartbeat Technology, SIL approved | E+H product page P 200 | Seul loop-powered (2-wire) — pas d'alimentation séparée nécessaire. Mais DN limité à 200 max et accuracy ±0.5% ± 2 mm/s (moins précis à bas débit que P 300)
Endress+Hauser | Promag P 300 | Advanced process | 8705 | — | OPTIFLUX 4300 | 316L, C22, Tantalum, Platinum, Titanium | PFA, PFA HT, PTFE | ±0.5% , ±0.2% (optional, Flat Spec) | ±0.1% | >1000:1 | ≥ 5 µS/cm | 0.01 to 10 m/s | -40°C | +180°C | PN 40 / Class 300 | 20K (JIS) | DN 15–600 | Pulsed DC | HART, PROFIBUS DP/PA, Modbus RS485, EtherNet/IP, PROFINET, Foundation Fieldbus | Yes | IP67 | Compact or remote, WLAN, web server | AC: 100–230 VAC / DC: 20–55 VDC | Heartbeat Technology, data logger | TI01224D | Full fieldbus + WLAN. Accuracy option ±0.2% vs Emerson ±0.15%. Mais pas de PFA+ ni de Class 600+
Endress+Hauser | Promag P 500 | Premium process | 9205 | — | OPTIFLUX 4400 | 316L, C22, Tantalum, Platinum, Titanium, Duplex | PFA, PFA HT, PTFE | ±0.5% , ±0.2% (optional, Flat Spec) | ±0.1% | >1000:1 | ≥ 5 µS/cm | 0.01 to 10 m/s | -40°C | +180°C | PN 40 / Class 300 | 20K (JIS) | DN 15–600 | Pulsed DC | HART, PROFIBUS DP/PA, Modbus RS485, EtherNet/IP, PROFINET, Foundation Fieldbus | Yes | IP67 | Remote only, up to 4 I/Os, WLAN, web server | AC: 100–230 VAC / DC: 20–55 VDC | Heartbeat Technology, data logger | TI01226D | Le plus complet E+H : Duplex 1.4462 (unique), 4 I/Os, remote only. Plus cher que 8705 mais plus de protocols
Krohne | Optiflux 1100 | Compact wafer | 8711 | Promag W 10 / P 10 | — | 316L, Hastelloy C, Titanium, Tantalum, Platinum | PTFE, PFA, ETFE | ±0.4% | ±0.1% | >1000:1 | ≥5 µS/cm (non-water ≥1 µS/cm) | 0.3–12 m/s | -25°C | +120°C | PN 40 (DN ≤50), PN 16 (DN 65–150) | Class 300 | DN 10–150 | Pulsed DC | HART, Modbus RS485 | Yes | IP67 | IFC 100 (compact or remote) | AC: 100–230 VAC / DC: 24 VDC | Conductivity, electrode error, ambient/process temperature | TD OPTIFLUX 1000 + TD IFC 100 | IFC 100 basique. Accuracy ±0.4% vs ±0.25% (8711) et ±0.5% (W 10). Moins précis mais plus de matériaux électrode que W 10
Krohne | Optiflux 1300 | Compact wafer, advanced | 8711 | Promag W 10 / P 10 | — | 316L, Hastelloy C, Titanium, Tantalum, Platinum | PFA (reinforced) | ±0.3% o.r. ± 1 mm/s | ±0.1% | 1000:1 | ≥ 5 µS/cm | 0.3–12 m/s | -25°C | +120°C | PN 40 (DN ≤50), PN 16 (DN 65–150) | Class 300 | PN 40 / Class 300 | Pulsed DC + Dual-frequency option | HART, Modbus RS485, PROFIBUS PA/DP, PROFINET, Foundation Fieldbus | Yes |  | IFC 300 (compact, field, wall, or rack mount) | AC: 100–230 VAC / DC: 24 VDC | Full suite + OPTICHECK + Virtual Reference + dual-frequency excitation | Krohne product page OPTIFLUX 1300 + TD IFC 300 | IFC 300 avec dual-frequency et Virtual Reference — ni 8711 ni W 10 n'offrent ça. Meilleur pour fluides difficiles en petit DN
Krohne | Optiflux 2100 | Standard water | 8750W | Promag W 400 | — | 316L, Hastelloy C, Titanium, Tantalum | PTFE, ETFE, Hard rubber, PU | ±0.3% | ±0.1% | >1000:1 | ≥5 µS/cm (water ≥20 µS/cm) | 0.3–12 m/s | -25°C | +120°C | PN 40 (DN ≤50), PN 16 (DN 65–150), PN 10 (DN 200+) | Class 300 | DN 25–1200 | Pulsed DC | HART, Modbus RS485 | Yes | IP67 / IP68 | IFC 100 (compact or remote) | AC: 100–230 VAC / DC: 24 VDC | Conductivity, electrode error, ambient/process temperature, OPTICHECK verification | TD OPTIFLUX 2000 + TD IFC 100 | Pression standard diminue avec DN (PN 10 au-delà de DN 200) vs 8750W (Class 150/300 constant). Plus de liners (4 vs 2)
Krohne | Optiflux 2050 | Basic water/wastewater | 8750W | Promag W 100 | — | 316L, Hastelloy C, Titanium, Tantalum | PTFE, ETFE, Hard rubber, PU | ±0.5% o.r. ± 2 mm/s | ±0.2% | 1000:1 | ≥5 µS/cm (water ≥20 µS/cm) | 0.3–12 m/s | -25°C | +120°C | PN 40 (DN ≤50), PN 16 (DN 65–150), PN 10 (DN 200+) | Class 300 | DN 25–1200 | Pulsed DC | HART, Modbus RS485 | Yes |  | IFC 050 (compact or remote) | AC: 100–230 VAC / DC: 24 VDC | Basic (conductivity, electrode error) + OPTICHECK | Krohne product page OPTIFLUX 2050 | Version budget IFC 050. Accuracy ±0.5% ± 2 mm/s (moins précis). Pas de dual-frequency. Tropicalized electronics pour humidité extrême
Krohne | Optiflux 2300 | Advanced water | 8750WA | Promag W 300/500 | — | 316L, Hastelloy C, Titanium, Tantalum | PTFE, ETFE, Hard rubber, PU | ±0.2% | ±0.1% | >1000:1 | ≥5 µS/cm (water ≥20 µS/cm) | 0.3–12 m/s | -25°C | +120°C | PN 40 (DN ≤50), PN 16 (DN 65–150), PN 10 (DN 200+) | Class 300 | DN 25–3000 | Pulsed DC + Dual-frequency option | HART, Modbus RS485, PROFIBUS PA/DP, PROFINET, Foundation Fieldbus | Yes | IP67 / IP68 | IFC 300 (compact, field, wall, or rack mount) | AC: 100–230 VAC / DC: 24 VDC | Full suite + OPTICHECK verification + Virtual Reference (no grounding rings needed) | TD OPTIFLUX 2000 + TD IFC 300 | Virtual Reference élimine les grounding rings — économie de coût et de maintenance vs 8750WA et W 300 qui en ont besoin
Krohne | Optiflux 4100 | Standard process | 8705 | Promag P 100 | — | 316L, Hastelloy C, Titanium, Tantalum, Platinum | PTFE, PFA, ETFE, PU, Hard rubber, Soft rubber | ±0.3% | ±0.1% | >1000:1 | ≥5 µS/cm (non-water ≥1 µS/cm) | 0.3–12 m/s | -25°C | +180°C | PN 40 (DN ≤80), PN 16 (DN 100–150), PN 10 (DN 200+) | Class 600 | DN 2.5–1200 | Pulsed DC | 4-20 mA HART, Modbus RS485 | Yes | IP67 / IP68 | IFC 100 (compact or remote) | AC: 100–230 VAC / DC: 24 VDC | Conductivity, electrode error, ambient/process temperature, OPTICHECK verification | TD OPTIFLUX 4000 + TD IFC 100 | Seul à offrir Soft rubber liner en standard — nécessaire pour rock slurry. 8705 et P 300 ne l'offrent pas en standard
Krohne | Optiflux 4300 | Advanced process | 8705 | Promag P 300 | — | 316L, Hastelloy C, Titanium, Tantalum, Platinum | PTFE, PFA, ETFE, PU, Hard rubber, Soft rubber | ±0.2% | ±0.1% | >1000:1 | ≥5 µS/cm (non-water ≥1 µS/cm) | 0.3–12 m/s | -25°C | +180°C | PN 40 (DN ≤80), PN 16 (DN 100–150), PN 10 (DN 200+) | Class 600 | DN 2.5–3000 | Pulsed DC + Dual-frequency option | HART, Modbus RS485, PROFIBUS PA/DP, PROFINET, Foundation Fieldbus | Yes | IP67 / IP68 | IFC 300 (compact, field, wall, or rack mount) | AC: 100–230 VAC / DC: 24 VDC | Full suite + OPTICHECK + Virtual Reference (no grounding rings) | TD OPTIFLUX 4000 + TD IFC 300 | Dual-frequency + Virtual Reference. Equivalent au P 300 en accuracy (±0.2%) mais avec suppression de bruit active pour slurries
Krohne | Optiflux 4400 | SIL-rated process | 9205 | Promag P 500 | — | 316L, Hastelloy C, Titanium, Tantalum, Platinum | PTFE, PFA, ETFE, PU, Hard rubber, Soft rubber | ±0.2% | ±0.1% | >1000:1 | ≥5 µS/cm (non-water ≥1 µS/cm) | 0.3–12 m/s | -25°C | +180°C | PN 40 (DN ≤80), PN 16 (DN 100–150), PN 10 (DN 200+) | Class 600 | DN 2.5–600 (SIL) / DN 2.5–2000 (non-SIL) | Pulsed DC | HART | Yes | IP67 / IP68 | IFC 400 (field or wall mount) | AC: 100–230 VAC / DC: 24 VDC | Detects gas bubbles, solids, electrode corrosion and deposits, conductivity changes, empty/partial pipe, disturbed flow profile, external magnetic fields + diagnostic test interval including fault response time of only 2 minutes | TD OPTIFLUX 4000 + TD IFC 400 | Seul modèle SIL 2/3 certifié parmi les 3 vendeurs. Fault response time 2 min. Ni Emerson ni E+H n'ont de SIL dédié pour magflow
Krohne | Optiflux 5100 | Aggressive/abrasive | 8707 | Promag P 500 | — | 316L, Hastelloy C, Titanium, Tantalum, Platinum | Ceramic (Al₂O₃) | ±0.3% | ±0.1% | >1000:1 | ≥5 µS/cm (non-water ≥1 µS/cm) | 0.3–12 m/s | -25°C | +180°C | PN 40 (DN ≤50) | Class 300 | DN 15–300 | Pulsed DC | HART, Modbus RS485 | Yes | IP67 / IP68 | IFC 100 (compact or remote) | AC: 100–230 VAC / DC: 24 VDC | Conductivity, electrode error, ambient/process temperature, OPTICHECK Electrodes: 316L SS, Hastelloy C, Titanium, Tantalum, Platinum | TD OPTIFLUX 5000 + TD IFC 100 | Ceramic tube — résistance abrasion extrême mais DN limité à 300 max. 8707 va jusqu'à DN 900 mais avec liner classique
Krohne | Optiflux 5300 | Advanced abrasive/aggressive with ceramic tube | 8707 | Promag P 300 | — | 316L, Hastelloy C, Titanium, Tantalum, Platinum | Ceramic (Al₂O₃) | ±0.15% o.r. ± 1 mm/s | ±0.1% | 1000:1 | ≥5 µS/cm (non-water ≥1 µS/cm) | 0.3–12 m/s | -25°C | +180°C | PN 40 (DN ≤50) | Class 300 | DN 2.5–300 (sandwich DN 2.5–100) | Pulsed DC + Dual-frequency option | HART, Modbus RS485, PROFIBUS PA/DP, PROFINET, Foundation Fieldbus | Yes | IP67 / IP69 | IFC 300 (compact, field, wall, or rack mount) | AC: 100–230 VAC / DC: 24 VDC | Full suite + OPTICHECK + Virtual Reference | Krohne product page OPTIFLUX 5300 | Accuracy ±0.15% — égal au 8705 optional, meilleur que P 300 (±0.2%). Ceramic + haute précision = le top pour slurries agressifs en petit DN

## Electrode Materials
Material | Emerson (Rosemount) | Endress+Hauser | Krohne | Availability | Temp Range (Acids) (°C) | Temp Range (Water) (°C) | Abrasion Resistance | Relative Cost | Source | Status | Notes
316L SS | ✓ | ✓ | ✗ | Standard | ✗ | up to 120°C | High | Very Low | JESA Internal DB — Ahmed Saidi 2021 | Vérifié | (missing some vendor verification)
Hastelloy C-276 | ✓ | ✓ | ✗ | Standard | 20–80°C | up to 120°C | Medium | Medium | JESA Internal DB — Ahmed Saidi 2021 | Vérifié | (E+H availability uncertain)
Hastelloy C-4 | ✓ | ✓ | ✓ | Standard | 20–80°C | up to 120°C | Medium | Medium | JESA Internal DB — Ahmed Saidi 2021 | Vérifié
Tantalum | ✓ | ✓ | ✓ | Standard | up to 120°C | up to 120°C | Low | Very High | JESA Internal DB — Ahmed Saidi 2021 | Vérifié
Platinum / Platinum-Iridium | ✓ | ✓ | ✓ | Standard | up to 120°C | up to 120°C | Low | Extremely High | JESA Internal DB — Ahmed Saidi 2021 | Vérifié
Titanium | ✓ | ✓ | ✓ | Standard | ✗ | up to 120°C | Medium | High | JESA Internal DB — Ahmed Saidi 2021 | Vérifié
Monel 400 | ✓ | ✓ | ✓ | Standard | ✗ (HF: 120°C) | up to 120°C | Medium | Medium | JESA Internal DB — Ahmed Saidi 2021 | Vérifié | (Ahmed Saidi says sea water, PDF says Titanium — discrepancy)
Alloy 20 | ✗ | ✗ | ✗ | Standard | up to 80°C | up to 120°C | Medium | Medium | Emerson PDS 00813-0100-4444 (listed but not standard at JESA) | À remplir | (vendor-only, not used at JESA)
Alloy C22 | ✓ | ✓ | ✓ | Standard | 20–80°C | up to 120°C | Medium | Medium | JESA Flowmeter Material Selection DB (PDF), Emerson PDS 00813-0100-4444, E+H TI01224D | Vérifié | (Krohne availability uncertain)
316 Ti | ✗ | ✗ | ✓ | Standard (Krohne only) | ✗ | up to 120°C | High | Low | Krohne TD OPTIFLUX 4000 | À remplir | (Krohne-only, not used at JESA)
Hastelloy B2 | ✗ | ✗ | ✓ | Standard | up to 120°C | ✗ (oxidizing) | Medium | Medium | Krohne TD OPTIFLUX 4000 | À remplir | (Krohne-only, not used at JESA)
Conductive Rubber | ✗ | ✓ | ✓ | Special order | ✗ | up to 80°C | Very High | Low | JESA Internal DB — Ahmed Saidi 2021 + JESA Lessons Learned (MIA/Daoui) | Vérifié
Hastelloy C | ✓ | ✓ | ✓ | Standard (Krohne only) | 20–80°C | up to 120°C | Medium | Medium | JESA Internal DB — Ahmed Saidi 2021 | Vérifié
Tungsten Carbide | ✗ | ✗ | ✓ | Special order | ✗ | up to 120°C | Extremely High | High | Krohne OPTIFLUX 5000 documentation | Vérifié | Mining applications only. Not chemically resistant — purely mechanical resistance.

## Liner Materials
Material | Emerson | Endress+Hauser | Krohne | Availability | Temp Range (°C) | Abrasion Resistance | Vacuum Suitable | Relative Cost | Source | Notes
PTFE | ✓ | ✓ | ✓ | Standard | -40 to +180°C | Low | No | Medium | App Ahmed Saidi 2021 (onglet "Mag netic database" + "Mag Flow material APP")+Emerson PDS + Krohne TD OPTIFLUX 4000 (section "Vacuum load") | Cannot be injection molded - weaker corrosion resistance than PFA at flange joints-DN range varies by vendor: E+H DN 15–600, Emerson DN 15–900, Krohne DN 2.5–1200
PFA | ✓ | ✓ | ✓ | Standard | -20 to +150°C | Low | Limited | Medium-High | App Ahmed Saidi 2021 (onglet "Mag netic database" + "Mag Flow material APP")+Emerson PDS + Krohne TD OPTIFLUX 4000 (section "Vacuum load") | Krohne PFA standard = 180°C grâce au renforcement SS mesh, équivalent au PFA HT de E+H-Default JESA liner for acids and most process fluids-DN limited: E+H DN 25–200, Krohne DN 2.5–150, Emerson DN 15–350 (PFA+)
PFA HT | ✗ | ✓ | ✗ | Standard (E+H only) | -20 to +180°C | Low | Limited | High | E+H TI01224D (Promag P 300), TI01225D (P 10), TI01226D (P 500) | Same PFA base - fluoropolymers have inherently low abrasion resistance-requires SS mesh reinforcement for partial vacuum-More expensive than standard PFA due to special high-temperature formulation
PFA+ | ✓ | ✗ | ✗ | Standard (Emerson only) | -20 to +177°C | Low | Yes (Limited) | High | Emerson PDS 00813-0100-4444 Rev AK | Same PFA base - mechanical improvement but no abrasion improvement-Better dimensional stability than standard PFA - resists vacuum and deformation better-More expensive than standard PFA due to reinforced construction + extended DN to 350
Hard Rubber (Ebonite) | ✓ | ✓ | ✓ | Standard | 0 to +80°C | Good | Yes | Low | App Ahmed Saidi 2021 + Emerson PDS | Not suitable for acids or solvents-Best value liner for water with particles
Soft Rubber (Natural) | ✓ | ✗ | ✓ | Standard | -20 to +60°C | Excellent | Yes | Low | App Ahmed Saidi 2021 | "soft rubber can support just 60°C"
Polyurethane | ✓ | ✓ | ✓ | Standard | -20 to +60°C | Excellent | No | Medium | App Ahmed Saidi 2021 +Emerson PDS 00813-0100-4444 | Good for abrasive but non-acidic fluids — alternative to soft rubber
Ceramic (Al2O3) | ✗ | ✗ | ✓ | Standard (Krohne only) | -20 to +180°C | Excellent | Yes | High | Krohne TD OPTIFLUX 5000 + page produit Krohne OPTIFLUX 5100 | Krohne only (OPTIFLUX 5100/5300)-DN limited to 300 max FDA compliant - also usable for hygienic applications
ETFE | ✓ | ✗ | ✓ | Standard | -20 to +120°C | Moderate | No | Medium | App Ahmed Saidi 2021 (onglet "Mag netic database" + "Mag Flow material APP") + Krohne TD OPTIFLUX 4000 + E+H TI01224D+Emerson PDS | Krohne offers ETFE for DN above 8" where PFA is not available
Polypropylene | ✗ | ✗ | ✗ | Not available | 0 to +80°C | Moderate | No | Low | ______ Propriété générale du matériau
PEX | ✓ | ✗ | ✗ | Standard (Emerson only) | -20 to +100°C | Moderate | No | Low | Emerson Rosemount 8700 PDS-Emerson Application Guide | Water only — no chemical resistance
Adiprene | ✓ | ✗ | ✗ | Special order (Emerson only) | -20 to +60°C | Excellent | No | Medium | ______ Propriété générale du matériau | intended for "oil and gas applications with high pressure, high salinity"
Neoprene | ✓ | ✗ | ✗ | Standard (Emerson only) | -20 to +80°C | Good | No | Low | Emerson Rosemount 8700 PDS-Emerson Application Guide | JESA standard for sea water — not resistant to acids
Linatex Rubber | ✓ | ✗ | ✗ | Standard (Emerson only) | -20 to +70°C | Excellent | No | Medium | Emerson Rosemount 8700 PDS-Emerson Application Guide | Alternative to soft rubber for slurries — longer lifespan but more expensive

## Drift Detection
Drift Indicator | What It Tells You | How to Measure | Baseline Method | Risk Level Rule | Root Cause | Corrective Action | Emerson Support | E+H Support | Krohne Support | Source | Notes
Electrode resistance change | Coating buildup or corrosion on electrodes | Ohm measurement via transmitter diagnostics | Set at commissioning, compare periodically | Abrasive fluid = High, Corrosive = Medium, Clean = Low | Fouling, scaling, chemical attack | Clean or replace electrodes | Yes | Yes | Yes | Emerson White Paper 88128, E+H SD03129DEN, Krohne OPTIFLUX 4400 product page | —Smart Meter Verification measures electrode circuit resistance against baseline. Rosemount 8705 offers field-replaceable electrodes(Emerson)— Heartbeat monitors E1/E2 electrode impedance continuously(E+H)—OPTICHECK measures electrode error. OPTIFLUX 4400 detects electrode corrosion and deposits automatically(Krohne)
Signal-to-noise ratio drop | Weak or noisy flow signal | Transmitter diagnostics / HART | Baseline noise level at stable flow | Slurry/multiphase = High, Acid = Medium, Water = Low | Electrode fouling, EMI, partial fill, slurry, two-phase flow | Check grounding, clean electrodes, increase damping, change coil drive frequency to 37.5 Hz (Emerson) | Yes | Yes | Yes | Emerson Manual 00809-0100-4727, E+H Heartbeat documentation, Krohne OPTIFLUX 4000 TD | — High Process Noise Detection (DA1 diagnostic suite), recommends increasing damping or switching coil drive frequency(Emerson)— Heartbeat Monitoring tracks noise trends over tim(E+H)—Dual-frequency excitation on OPTIFLUX 4300/4400 actively suppresses process noise(Krohne)
Zero drift | Measurement reads non-zero at zero flow | Zero-point verification test with block valves closed, full pipe, no flow | Zero-point recorded at commissioning | Same for all fluids — always Low | Residue on electrodes, installation issue, EMI | Re-zero via auto-zero function, clean electrodes | Yes | Yes | Yes | Emerson Manual 00809-0100-4727 (Section 4), E+H BA01063DEN, Krohne IFC 100/300 TD | — Auto-zero procedure completes automatically in approximately 90 seconds(Emerson)— Auto-zero function available, Heartbeat Verification checks zero stability(E+H)— Auto-zero available on all IFC converters(Krohne)
Empty pipe false readings | Meter reads flow when pipe is empty/partial | Electrode conductivity check via EPD (Empty Pipe Detection) function | EPD configured at commissioning | Gravity drain/pump shutdown applications = Medium, others = Low | Partial fill, air entrainment, pump shutdown, gravity drain | Check process conditions, adjust EPD settings, verify installation position | Yes | Yes | Yes | Emerson PDS 00813-0100-4444, E+H TI01580D (W 10), Krohne OPTIFLUX 4400 product page | — Tunable Empty Pipe Detection, standard on all models. Configurable sensitivity(Emerson)— Dedicated EPD electrode, impedance monitored via Heartbeat. Can assign to switch output(E+H)— Empty pipe detection on all OPTIFLUX models. OPTIFLUX 4400 detects partially filled pipe automatically(Krohne)
Coil insulation degradation | Moisture in coil housing, aging insulation, coil circuit failure | Coil resistance measurement + coil current monitoring via transmitter | Coil resistance 50–240 Ohm (E+H ref), coil current ±0.1% | Age >15 years = Medium, moisture environment = High, others = Low | Moisture ingress, vibration, aging, environmental exposure | Replace coil assembly, check IP rating integrity, verify cable glands | Yes | Yes | Partial | Emerson White Paper 88128, E+H SD03129DEN (Heartbeat Verification report with actual values), Krohne IFC 100 TD | — Coil Fault diagnostic (standard). Smart Meter Verification checks coil circuit resistance and coil signature (magnetic field strength) against baseline(Emerson)— Heartbeat monitors coil resistance (example range: 50–240 Ohm), shot time symmetry (0.9–1.1), coil current offset (±0.1%), and coil current loss deviation (±10%)(E+H)— Coil circuit monitoring available, detected via OPTICHECK service tool(Krohne)
Liner integrity loss | Chemical attack or abrasion on liner | Visual inspection during scheduled maintenance / indirectly via process noise increase and measurement instability | No direct baseline — monitor via process noise | Abrasive = High, Corrosive = Medium, Clean = Low | Chemical attack, abrasion (especially slurries), delamination, vacuum damage | Replace sensor. Preventive: select correct liner for fluid (soft rubber for rock slurry, PFA for acids) | Partial | Partial | Partial | JESA Lessons Learned (MIA/Daoui: rock phosphate slurry destroyed liner and electrode), Emerson Manual 00809-0100-4727 | — No direct liner diagnostic. Detected indirectly through increased process noise or electrode signal instability(Emerson)— Heartbeat Monitoring build-up index can indicate liner surface changes(E+H)— Detected indirectly via electrode and conductivity diagnostics(Krohne)
Process noise increase | Unstable or noisy reading | Standard deviation of measurement | Baseline noise at stable operating conditions | Slurry = High, Two-phase = High, Clean = Low | Two-phase flow, cavitation, slurry | Check process, resize meter | Yes | Yes — Heartbeat Monitoring continuously tracks noise levels and trends over time. Noise increase triggers warning alarm. Historical trending helps predict liner degradation before failure | Yes — Dual-frequency excitation on IFC 300/400 actively suppresses process noise by combining low-frequency signal (accuracy) with high-frequency signal (noise rejection). OPTIFLUX 4400 automatically detects gas bubbles and solids content changes | Emerson Manual 00809-0100-4727, Krohne OPTIFLUX 4400 product page | Max velocity 3 m/s recommended for slurries (App 2024), 6 m/s absolute maximum (JESA Lessons Learned). High velocity = more noise = faster liner wear. Emerson 8707 uses High-Signal excitation (10x coil current) as alternative to dual-frequency.
Grounding fault | Erratic readings, noise | Ground resistance measurement | Ground impedance <1 Ohm at commissioning | FRP/PVC pipe = Medium, Metal pipe = Low | Corroded grounding rings/electrodes | Restore/replace grounding | Yes | Yes — Heartbeat Verification checks grounding impedance. Must be <1 Ohm per E+H specification. Detected during commissioning and periodic verification | Yes — Virtual Reference on IFC 300 eliminates the need for grounding rings entirely (non-SIL applications only). On IFC 400 (SIL), physical grounding remains mandatory. OPTIFLUX 4400 detects external magnetic fields that may indicate grounding issues | JESA Lessons Learned Note 6, E+H BA01063DEN, Krohne IFC 300 TD | Non-conductive pipes (FRP, PVC, HDPE, rubber-lined) always require grounding rings. Conductive pipes (SS, CS) need grounding straps only. Grounding ring material must match fluid: Tantalum for acids, Monel 400 for sea water, Hastelloy C-4 for mineral slurries.
Span drift | Gradual loss of accuracy over full range | Calibration verification | Factory calibration stored in transmitter | All fluids = Low (gradual, detectable via SMV/Heartbeat) | Electrode degradation, EMI | Recalibrate | Partial | Yes — Heartbeat Verification confirms meter is operating within specified measuring tolerance over its entire useful lifetime. Generates traceable verification report (ISO 17025 compatible) that can replace wet calibration in some applications. Detects span shift through coil and electrode parameter trending | Partial — OPTICHECK service tool provides on-site verification of calibration status. Long-term stability specified at ±0.1% of measured value. IFC 300/400 store factory calibration data for comparison. No continuous span monitoring like Heartbeat | Emerson White Paper 88128, E+H SD03129DEN, Krohne TD OPTIFLUX 4000 | Span drift is caused by coil aging (magnetic field weakening), electrode surface changes, or EMI. Emerson detects via Coil Signature (measures magnetic field strength — any change = calibration shift). If span drift exceeds ±0.5%, recalibration required. If coil signature has shifted, sensor may need replacement.
Converter self-test failure | Electronics malfunction | Built-in diagnostics | All fluids = Low (gradual, detectable via SMV/Heartbeat) | Same for all — always Low unless age >15 years | Component failure, power issues | Replace converter | Yes | Yes — Heartbeat Diagnostics monitors all relevant supply voltages, overvoltage protection, low-side current, shield driver circuit, and overall system condition. Detects component degradation before complete failure. Alerts via HART/PROFIBUS/Ethernet with NAMUR NE 107 compliant status codes (failure, function check, out of specification, maintenance required) | Yes — OPTIFLUX 4400 (IFC 400) has the most advanced self-test with diagnostic test interval including fault response time of only 2 minutes. Detects external magnetic fields, disturbed flow profiles, and electronics failures simultaneously. IFC 100/300 have basic continuous self-test with longer response time | Emerson PDS 00813-0100-4444 (Table 2), E+H SD01847DEN (Promag 400 Heartbeat), Krohne OPTIFLUX 4400 product page | All 3 vendors provide continuous self-test as standard on all models — no additional configuration needed. Common root causes: power supply instability (voltage spikes or drops), extreme ambient temperature (>60°C or <-20°C), moisture ingress into terminal compartment (IP rating compromised), component aging (typically after 15+ years). If self-test fails: first check power supply stability, then check ambient conditions, then check cable gland tightness. If persistent → replace converter/transmitter.
Temperature deviation | Process temperature exceeds liner rating or deviates from design | Process temperature sensor + transmitter ambient temperature monitoring | Design temperature from process datasheet | Near liner limit = High, >20°C margin = Low | Process upset, heat exchanger failure, ambient conditions | Verify process temperature, check liner temperature rating | Yes — Electronics Temperature monitoring | Yes — Heartbeat monitors ambient and process temperature | Yes — OPTIFLUX 4400 ambient/process temperature monitoring |  | PFA liner fails above 150°C (standard) or 180°C (HT). Soft rubber fails above 60°C. Temperature excursion = liner damage risk
Flow profile disturbance | Upstream piping changes or valve position affects measurement accuracy | Signal symmetry check between electrodes | Symmetric signal at commissioning | <5D upstream = High, ≥5D = Low | Insufficient straight run upstream (<5D), partially open valve, pump pulsation | Verify installation meets 5D upstream / 3D downstream requirement. Check valve positions | Partial — detected via process noise increase | Partial — Heartbeat detects asymmetric flow | Yes — OPTIFLUX 4400 explicitly detects disturbed flow profiles |  | Common issue at OCP plants where piping is congested. Affects accuracy without triggering other alarms
Conductivity change | Process fluid composition has changed or contamination detected | Conductivity measurement via transmitter (built-in on all magflows) | Baseline conductivity at normal operation | Demineralized water = High, Acids = Medium, others = Low | Process upset, dilution, concentration change, contamination | Verify process conditions, check upstream operations | Yes — continuous conductivity monitoring | Yes — Heartbeat Monitoring tracks conductivity trends | Yes — OPTIFLUX 4400 detects conductivity changes automatically |  | Critical for demineralized water (can drop below 5 µS/cm) and for acids (concentration change = different electrode requirement)

## Fluid-Material Matrix
Service Fluid | Fluid type | Coil Housing Material | Flow Tube Material | Liner Material | Electrode Material | Flange Material | Remote Transmitter Material | 2" Mounting Bracket Material | Penetrant Ring (Abrasive fluids) | Grounding Ring (Non-conductive pipe) | O-Ring Material | Flange Coating | Pipe Material | Remarks
100% (w/w) Ester solution | Corrosive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Hastelloy C-4 | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | EPDM | None | Metal
5% Phosphoric Ester solution | Corrosive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Hastelloy C-4 | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | EPDM | None | Metal
Acidic process water | Corrosive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Tantalum | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | 2x Tantalum grounding rings | Viton | PTFE coating | FRP | Acids on FRP → 2x Tantalum grounding rings required
Amine solution | Corrosive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Hastelloy C-4 | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | EPDM | None | Metal
Cake wash water | Clean | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Hastelloy C | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | EPDM | None | Metal
Caustic solution | Corrosive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Titanium ou Platinum | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | Viton | None | Metal | Titanium: High cost, resistant to alkalis. Platinum: Extremely High cost, universal chemical resistance
Cloth wash water | Clean | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Hastelloy C / Platinum | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | EPDM | None | Metal | Hastelloy C: Medium cost, sufficient for water. Platinum: Extremely High cost, no added benefit for this fluid
Cooling water supply | Clean | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Hastelloy C / Platinum | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | EPDM | None | Metal | Hastelloy C: Medium cost, sufficient for water. Platinum: Extremely High cost, no added benefit for this fluid
sea water based mineral Froth | Corrosive + Abrasive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Hastelloy C-4 | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | Oui | 2x Hastelloy C-4 grounding rings | Viton | PTFE coating | FRP | Abrasif + corrosif + non-conductive pipe → penetrant ring + 2x Hastelloy C-4 grounding rings
sea water based mineral Slurry | Corrosive + Abrasive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Hastelloy C-4 | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | Oui | 2x Hastelloy C-4 grounding rings | Viton | PTFE coating | FRP | Abrasif + corrosif + non-conductive pipe → penetrant ring + 2x Hastelloy C-4 grounding rings
Defoamer | Clean | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Platinum | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | EPDM | None | Metal
Ester solution | Clean | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Hastelloy C-4 | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | EPDM | None | Metal | Standard clean service
Filtered water | Clean | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Alloy C22 | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | EPDM | None | Metal
FSA (Fluorosilicic acid) | Corrosive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Platinum | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | Viton | PTFE coating | Metal
High pressure Gland seal water | Clean | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Hastelloy C | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | EPDM | None | Metal
Low chloride water | Clean / Mild Corrosive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Hastelloy C / Platinum | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | EPDM | None | Metal | Hastelloy C: Medium cost, sufficient for water. Platinum: Extremely High cost, no added benefit for this fluid
Low pressure Gland seal water | Clean | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Hastelloy C | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | EPDM | None | Metal
Medium pressure Gland seal water | Clean | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Hastelloy C | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | EPDM | None | Metal
Micronutrient solution | Mild Corrosive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Platinum | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | Oui | N/A | Viton | PTFE coating | Metal | Fluide abrasif → penetrant ring requis
Mineral Froth | Abrasive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | Polyurethane / Hard Rubber | Hastelloy C-4 / Tungsten Carbide | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | Oui | 2x Hastelloy C-4 grounding rings | Viton | PTFE coating | FRP / Metal | High abrasion – use reinforced liner + hard electrodes
Mineral Slurry | Abrasive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | Polyurethane / Linatex Rubber | Hastelloy C-4 / Tungsten Carbide | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | Oui | 2x Hastelloy C-4 grounding rings | Viton | PTFE coating | FRP / Metal | Severe abrasion – prioritize predictive monitoring for liner wear
Potable water | Clean | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Alloy C22 | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | EPDM | None | Metal
Pressurized sea water | Abrasive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Hastelloy C-4 | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | Oui | 2x Hastelloy C-4 grounding rings | Viton | PTFE coating | FRP | Abrasif + non-conductive pipe → penetrant ring + 2x Hastelloy C-4 grounding rings
Process condensate | Clean | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Alloy C22 | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | EPDM | None | Metal
Process water | Clean | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Hastelloy C-4 | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | EPDM | None | Metal
Product Acid | Corrosive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Platinum | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | Viton | PTFE coating | Metal
Raw water | Clean | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Alloy C22 | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | EPDM | None | Metal
Reactor Slurry | Abrasive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Platinum / Hastelloy / 80%Pt-20%Ir | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | Oui | Grounding strap | Viton | PTFE coating | Metal | 3 options d'électrode selon disponibilité vendeur
Recovered Process Water | Clean | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Hastelloy C-4 | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | Oui | 2x Hastelloy C-4 grounding rings | EPDM | None | FRP | Abrasif + non-conductive pipe → penetrant ring + 2x Hastelloy C-4 grounding rings
Recovered Liquor | Corrosive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Platinum | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | Viton | PTFE coating | Metal
RO water | Clean | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Hastelloy C-4 | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | EPDM | None | Metal
Rock slurry | Abrasive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | Soft rubber | Conductive rubber / Hastelloy / 80%Pt-20%Ir | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | Oui | N/A | Viton | PTFE coating | Metal | MIA/Daoui: conductive rubber = moins de bruit que métal. 3 options car conductive rubber pas disponible chez tous les vendeurs
Scrubber liquor | Corrosive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Platinum | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | Viton | PTFE coating | Metal
Sea water | Corrosive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | Neoprene ou PFA | Monel 400 ou Titanium | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | 2x Monel 400 grounding rings | EPDM | None | FRP | Monel 400 + PFA: better chemical resistance, higher temp range. Titanium + Neoprene: lower cost, limited to 80°C
Slurry | Abrasive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Platinum / Hastelloy / 80%Pt-20%Ir | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | Oui | N/A | Viton | PTFE coating | Metal | Abrasif → penetrant ring requis. 3 options d'électrode
Strong phosphoric acid | Corrosive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Tantalum | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | 2x Tantalum grounding rings | Viton | PTFE coating | FRP | Acide sur FRP → 2x Tantalum grounding rings
Sulfuric Acid | Corrosive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Tantalum / Platinum / 80%Pt-20%Ir | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | 2x Tantalum grounding rings | Viton | PTFE coating | FRP | Acide sur FRP → 2x Tantalum grounding rings. 3 options d'électrode
Very High pressure Gland seal water | Clean | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Hastelloy C | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | EPDM | None | Metal
Waste disposal | Corrosive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Platinum | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | EPDM | None | Metal
Wash water | Clean | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Hastelloy C | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | EPDM | None | Metal
Weak phosphoric acid | Corrosive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Tantalum | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | Viton | PTFE coating | Metal | Can be on FRP → then Tantalum grounding rings needed
Gypsum Slurry | Abrasive + Corrosive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Platinum / Hastelloy C-4 | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | Oui | Grounding strap | Viton | PTFE coating | Metal | Major OCP byproduct from phosphoric acid production. Very abrasive. Max velocity should follow the 3 m/s recommended rule, not 6 m/s.
Lime Slurry | Abrasive + Mild Corrosive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Platinum / Hastelloy C-4 | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | Oui | Grounding strap | Viton | PTFE coating | Metal | Alkaline and abrasive. Unlike acid slurries, Titanium works here because Titanium resists alkalis. It's a third electrode option not available for acid slurries.
Dilute Phosphoric Acid | Corrosive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Tantalum | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | 2x Tantalum grounding rings | Viton | PTFE coating | FRP
Waste Water | Corrosive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Hastelloy C / Alloy C22 | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | EPDM | None | Metal
Demineralized Water | Clean | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Alloy C22 | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | EPDM | None | Metal | This is the riskiest one. Demineralized water has very low conductivity (can be <1 µS/cm after full treatment). Magflows require ≥5 µS/cm.
High Chloride Water | Corrosive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Hastelloy C-4 / Titanium | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | Viton | None | Metal
Scrubber Liquor Slurry | Abrasive + Corrosive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Platinum | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | Oui | Grounding strap | Viton | PTFE coating | Metal
Treatment Chemical Phosphate | Corrosive | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Tantalum / Platinum | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | Viton | PTFE coating | Metal
Gland seal water LP/MP/HP/VHP | Clean | SS 304L CS with corrosion resistant paint Die cast alu with corrosion resistant paint | SS316L | PFA / PTFE | Hastelloy C / 316L | SS316L | Die Cast Alu with Epoxy Coated | SS 316 | N/A | Grounding strap | EPDM | None | Metal | Clean water service – standard configuration

## Sources & Evolution
Category | Item | Details | URL / Access | Update Frequency | Notes | Status
Vendor Portal | Emerson Rosemount Mag Flow | Product specs, selection guides, sizing tools | emerson.com | Quarterly
Vendor Portal | Endress+Hauser Promag | Product configurator, technical documentation | endress.com | Quarterly
Vendor Portal | Krohne OPTIFLUX | Product data, application reports | krohne.com | Quarterly
Standards | ISA S5.1 | Instrumentation symbology | isa.org | As revised
Standards | IEC 61298 | Flowmeter testing standards | iec.ch | As revised
Internal | JESA Document Server | P&IDs, datasheets, project specs | Internal network | Per project
Internal | OCP Maintenance Records | Failure logs, calibration history, drift data | Internal | Continuous | Key for drift detection pillar
MATERIAL EVOLUTION
New Material | Ceramic liners (Al2O3) | Extreme abrasion resistance for mining/slurry |  | ~2000s | Offered by Krohne, E+H
New Material | Diamond-like carbon (DLC) electrode coating | Anti-fouling, anti-scaling properties |  | ~2015s | Emerging technology
New Technology | Dual-frequency excitation | Better noise immunity in slurry applications |  | ~2010s | Krohne, E+H
New Material | Advanced alloys (Alloy 22) | Wider corrosion resistance than Hastelloy C-276 |  | ~2010s | Emerson
JESA Internal DB | Ahmed Saidi | Flow meter technology & material selection V2021 | Internal network, Per project
JESA | Hicham, Abderrahman, FatimaElhadi 2020 | Sizing Application | Internal network, Per project
JESA | PDF | Flowmeter Material Selection DB (PDF) | Internal network, Per project
JESA
Class ASME | Pression à 20°C (bar) | Pression à 100°C (bar) | Pression à 200°C (bar)
150 | 19.6 | 17.7 | 15.3
300 | 51 | 46.6 | 40.4
600 | 102 | 93.2 | 80.8
900 | 153 | 139.8 | 121.2
1500 | 255 | 233 | 202
2500 | 425 | 388.3 | 336.7
PN (DIN) | Pression à 20°C (bar)
PN 6 | 6
PN 10 | 10
PN 16 | 16
PN 25 | 25
PN 40 | 40
RÈGLES GÉNÉRALES JESA (applicables à tous les fluides)
1. Pour réduire les coûts, on peut utiliser SS304L pour le flow tube et la bride avec accord client (au lieu de SS316L).
2. Pour les petits DN : PFA avec SS mesh. Pour les grands DN : PTFE/ETFE liners.
3. Soft rubber est recommandé pour le rock slurry (basse température) car PFA/PTFE seront endommagés par la boue.
4. Soft rubber supporte max 60°C, PTFE et PFA supportent 120°C, PFA haute température supporte 180°C → la température est un critère de sélection du liner.
5. Si la tuyauterie est métallique (conductrice) → grounding strap suffit. Si FRP (non-conductrice) → grounding ring obligatoire.
6. Le matériau du grounding ring doit être compatible avec le fluide (Tantalum pour acides, Monel 400 pour eau de mer, Hastelloy C-4 pour slurries minéraux).
7. Les fluides abrasifs (slurries, froth, micronutrient) nécessitent des penetrant rings pour protéger le liner.
8. Vitesse max : 6 m/s pour les slurries, 10 m/s pour les liquides.
SOURCE
• JESA Internal DB — Ahmed Saidi, Flow meter technology & material selection V2021 (onglet 'Mag netic database')
• JESA Flowmeter Material Selection DB (PDF) — Flow_meters_material_selection_data_base__Final_revision_.pdf
• JESA Lessons Learned — Notes & Lessons Learned (page 9 du PDF), incluant MIA/Daoui
• Q37092-00-IC-DAS-00017 — PMP CAP Plant project datasheets (14 magnetic flowmeters)

## CAPEX Components
A. ELECTRODE MATERIALS — Relative Cost
Material | Relative Cost | Cost Score (1-6) | Source
316L SS | Very Low | 1 | JESA Internal DB — Ahmed Saidi 2021
Conductive Rubber | Low | 2 | JESA Lessons Learned (MIA/Daoui)
Alloy C22 | Medium | 3 | JESA Internal DB — Ahmed Saidi 2021
Hastelloy C-276 | Medium | 3 | JESA Internal DB — Ahmed Saidi 2021
Hastelloy C-4 | Medium | 3 | JESA Internal DB — Ahmed Saidi 2021
Hastelloy C | Medium | 3 | JESA Internal DB — Ahmed Saidi 2021
Monel 400 | Medium | 3 | JESA Internal DB — Ahmed Saidi 2021
Alloy 20 | Medium | 3 | Emerson PDS 00813-0100-4444
Hastelloy B2 | Medium | 3 | Krohne TD OPTIFLUX 4000
316 Ti | Low | 2 | Krohne TD OPTIFLUX 4000
Titanium | High | 4 | JESA Internal DB — Ahmed Saidi 2021
Tantalum | Very High | 5 | JESA Internal DB — Ahmed Saidi 2021
Platinum / Pt-Ir | Extremely High | 6 | JESA Internal DB — Ahmed Saidi 2021
B. LINER MATERIALS — Relative Cost
Material | Relative Cost | Cost Score (1-5) | Source
Hard Rubber (Ebonite) | Low | 1 | Emerson PDS 00813-0100-4444
Soft Rubber (Natural) | Low | 1 | Emerson PDS 00813-0100-4444
Neoprene | Low | 1 | Emerson PDS 00813-0100-4444
PEX | Low | 1 | Emerson PDS 00813-0100-4444
Polyurethane (PU) | Medium | 2 | Emerson PDS 00813-0100-4444
ETFE | Medium | 2 | E+H TI01224D, Krohne TD OPTIFLUX 4000
Linatex Rubber | Medium | 2 | Emerson PDS 00813-0100-4444
PTFE | Medium | 2 | Emerson / E+H / Krohne datasheets
Adiprene | Medium | 2 | Emerson PDS 00813-0100-4444
PFA | Medium-High | 3 | Emerson / E+H / Krohne datasheets
PFA HT (E+H) | High | 4 | E+H TI01224D
PFA+ (Emerson) | High | 4 | Emerson PDS 00813-0100-4444
Ceramic (Al2O3) | High | 4 | Krohne TD OPTIFLUX 5000
C. GROUNDING & PROTECTION — Relative Cost
Component | Relative Cost | Cost Score (1-5) | Source
Grounding strap (conductive pipe) | Very Low | 1 | JESA Internal DB — Ahmed Saidi 2021
O-Ring EPDM | Very Low | 1 | Q37092-00-IC-DAS-00017 (PMP CAP)
O-Ring Viton | Low | 1 | Q37092-00-IC-DAS-00017 (PMP CAP)
Flange coating PTFE | Low | 1 | Q37092-00-IC-DAS-00017 (PMP CAP)
Grounding ring 904L SS (x2) | Medium | 2 | Q37092-00-IC-DAS-00017 (PMP CAP)
Grounding ring Hastelloy C-4 (x2) | Medium | 3 | JESA Internal DB — Ahmed Saidi 2021
Grounding ring Monel 400 (x2) | Medium | 3 | JESA Internal DB — Ahmed Saidi 2021
Penetrant ring Hastelloy C-276 | Medium | 3 | JESA Flowmeter Material Selection DB (PDF)
Grounding ring Tantalum (x2) | High | 4 | JESA Flowmeter Material Selection DB (PDF)
D. VENDOR MODELS (Capteur + Transmetteur) — Relative Cost
Vendor — Model | Relative Cost | Cost Score (1-5) | Tier
Emerson — 8711 (Wafer) | Low | 1 | Compact/Budget
E+H — Promag W 10 | Low | 1 | Compact/Budget
E+H — Promag P 10 | Low | 1 | Compact/Budget
Krohne — OPTIFLUX 1100 | Low | 1 | Compact/Budget
Emerson — 8750W | Medium | 2 | Standard Water
Krohne — OPTIFLUX 2100 | Medium | 2 | Standard Water
E+H — Promag W 400 | Medium | 2 | Standard Water
Krohne — OPTIFLUX 4100 | Medium | 3 | Standard Process
Krohne — OPTIFLUX 5100 | Medium | 3 | Standard Abrasive
Emerson — 8705 | High | 3 | Advanced Process
Emerson — 8707/Slurry | High | 3 | Advanced Slurry
E+H — Promag P 300 | High | 3 | Advanced Process
Krohne — OPTIFLUX 2300 | High | 3 | Advanced Water
Krohne — OPTIFLUX 4300 | High | 4 | Advanced Process
E+H — Promag W 500 | High | 4 | Premium Water
E+H — Promag P 500 | Very High | 4 | Premium Process
Emerson — 9205+9732 | Very High | 4 | Premium Next-Gen
Krohne — OPTIFLUX 4400 | Very High | 5 | SIL-Certified

## OPEX Estimation
A. CALIBRATION — Fréquence et coût par type de fluide
Type de fluide | Fréquence calibration (mois) | Temps arrêt (heures) | Réduction avec SMV/Heartbeat/OPTICHECK | Fréquence étendue (mois) | Source
Eau potable / Raw water | 24 | 2-4h | -50% | 48 | Benchmark industrie
Filtered water / Process water | 24 | 2-4h | -50% | 48 | Benchmark industrie
Weak phosphoric acid | 12 | 2-4h | -50% | 24 | Benchmark industrie
Strong phosphoric acid | 12 | 2-4h | -50% | 24 | Benchmark industrie
98% Sulfuric acid | 12 | 2-4h | -50% | 24 | Benchmark industrie
TSP / Ammonium phosphate slurry | 6 | 4-6h | -30% | 9 | Benchmark industrie
Rock slurry | 6 | 4-6h | -30% | 9 | Benchmark industrie + MIA/Daoui
Sea water | 12 | 2-4h | -50% | 24 | Benchmark industrie
Caustic solution | 12 | 2-4h | -50% | 24 | Benchmark industrie
Mineral froth / mineral slurry | 6 | 4-6h | -30% | 9 | Benchmark industrie
B. DURÉE DE VIE ESTIMÉE DES COMPOSANTS (années)
Composant | Eau / Water services | Acides / Acid services | Slurries / Abrasive | Sea water | Source
Liner PFA | 15-20 ans | 10-15 ans | 5-8 ans | 12-15 ans | Benchmark industrie
Liner PTFE | 15-20 ans | 8-12 ans | 3-5 ans | 10-12 ans | Benchmark industrie
Liner Soft Rubber | N/A | N/A | 3-5 ans | N/A | JESA Lessons Learned
Liner Ceramic | N/A | 15-20 ans | 10-15 ans | N/A | Krohne documentation
Liner Neoprene | N/A | N/A | N/A | 8-12 ans | Emerson documentation
Electrode Alloy C22 | 15-20 ans | N/A | N/A | N/A | Benchmark industrie
Electrode Tantalum | 20+ ans | 15-20 ans | N/A | N/A | Benchmark industrie
Electrode Platinum | 20+ ans | 20+ ans | 10-15 ans | 20+ ans | Benchmark industrie
Electrode Titanium | 20+ ans | N/A | N/A | 15-20 ans | Benchmark industrie
Electrode Conductive Rubber | N/A | N/A | 3-5 ans | N/A | JESA Lessons Learned (MIA/Daoui)
Electrode Hastelloy C-4 | 15-20 ans | 10-15 ans | 8-12 ans | 12-15 ans | Benchmark industrie
O-Ring Viton | 5-8 ans | 3-5 ans | 3-5 ans | 5-8 ans | Benchmark industrie
O-Ring EPDM | 5-8 ans | N/A | N/A | 3-5 ans | Benchmark industrie
Bobines (Coils) | 20+ ans | 20+ ans | 20+ ans | 20+ ans | Vendor documentation
Transmetteur/Converter | 15-20 ans | 15-20 ans | 15-20 ans | 15-20 ans | Vendor documentation
Grounding ring | 10-15 ans | 8-12 ans | 5-8 ans | 8-12 ans | Benchmark industrie
C. ÉNERGIE — Consommation par modèle
Vendor — Model | Power Supply | Consumption (W) | Annual kWh (24/7 operation) | Relative Cost (20 ans) | Source
Emerson - 8705 + 8732 | AC: 90-250 VAC / DC: 12-42 VDC | 20 W | 175 kWh/an | Very Low | PDS 00813-0100-4444
Emerson - 8750W | AC: 90-250 VAC / DC: 12-42 VDC | 20 W | 175 kWh/an | Very Low | PDS 00813-0300-4750
Emerson - 8711 / 8707 | AC: 90-250V / DC: 12-42V | 20 W | 175 kWh/an | Very Low | PDS 00813-0100-4444
Emerson - 9205 | AC: 90-250V / DC: 12-42V | 20 W | 175 kWh/an | Very Low | PDS 00813-0200-4750
E+H - Promag W 10/100 | DC: 18-30 VDC | 15 W | 131 kWh/an | Very Low | TI01580D
E+H - Promag P 10 / 100 | DC: 18-30 VDC | 15 W | 131 kWh/an | Very Low | TI01580D
E+H - Promag P 200 | DC: 18-30 VDC | 5 W | 44 kWh/an | Very Low | TI01223D (2-wire)
E+H - Promag P 300 | AC: 100-230 VAC / DC: 20-55 VDC | 20 W | 175 kWh/an | Very Low | TI01224D
E+H - Promag W/P 400/500 | AC: 100-230 VAC / DC: 20-55 VDC | 25 W | 219 kWh/an | Very Low | TI01226D
Krohne - OPTIFLUX + IFC 050 | AC: 100-230V / DC: 24V | 12 W | 105 kWh/an | Very Low | TD IFC 050
Krohne - OPTIFLUX + IFC 100 | AC: 100-230 VAC / DC: 24 VDC | 18 W | 158 kWh/an | Very Low | TD IFC 100
Krohne - OPTIFLUX + IFC 300 | AC: 100-230 VAC / DC: 24 VDC | 22 W | 193 kWh/an | Very Low | TD IFC 300
Krohne - OPTIFLUX + IFC 400 | AC: 100-230 VAC / DC: 24 VDC | 22 W | 193 kWh/an | Very Low | TD IFC 400

## TCO Scoring Examples
TCO SCORING — Exemples de configurations complètes JESA
Configuration | Fluid | Model (Score) | Electrode (Score) | Liner (Score) | Grounding (Score) | Accessories (Score) | TOTAL CAPEX Score
Config 1 — Water budget | Potable water | OPTIFLUX 2100 (2) | Alloy C22 (3) | PTFE (2) | Strap (1) | EPDM + None (1) | 0
Config 2 — Water premium | Potable water | Promag W 500 (4) | Alloy C22 (3) | Hard rubber (1) | Strap (1) | EPDM + None (1) | 0
Config 3 — Acid standard | Strong H3PO4 | OPTIFLUX 4100 (3) | Tantalum (5) | PFA (3) | 2x Ta rings (4) | Viton + PTFE coat (1) | 0
Config 4 — Acid premium | Strong H3PO4 | 8705 (3) | Tantalum (5) | PFA+ (4) | 2x Ta rings (4) | Viton + PTFE coat (1) | 0
Config 5 — Slurry Emerson | Rock slurry | 8707 Slurry (3) | Cond. rubber (2) | Soft rubber (1) | N/A (0) | Viton + PTFE coat + PR (2) | 0
Config 6 — Slurry Krohne | Rock slurry | OPTIFLUX 5100 (3) | Platinum (6) | Ceramic (4) | N/A (0) | Viton + PTFE coat + PR (2) | 0
Config 7 — Sea water std | Sea water | OPTIFLUX 4100 (3) | Monel 400 (3) | PFA (3) | 2x Monel rings (3) | EPDM + None (1) | 0
Config 8 — Sea water Emerson | Sea water | 8705 (3) | Titanium (4) | Neoprene (1) | 2x Monel rings (3) | EPDM + None (1) | 0
Config 9 — Mineral slurry | Mineral froth | OPTIFLUX 4300 (4) | Hast. C-4 (3) | PFA (3) | 2x Hast C-4 (3) | Viton + PTFE + PR (2) | 0
Config 10 — SIL acid | 98% H2SO4 | OPTIFLUX 4400 (5) | Tantalum (5) | PFA (3) | 2x Ta rings (4) | Viton + PTFE coat (1) | 0
Config 11 — FSA | FSA | Promag P 300 (3) | Platinum (6) | PFA (3) | Strap (1) | Viton + PTFE coat (1) | 0
Config 12 — Gland seal | HP Gland seal water | Promag W 400 (2) | Hast. C (3) | PTFE (2) | Strap (1) | EPDM + None (1) | 0
GUIDE D'INTERPRÉTATION DU SCORE TCO
Score 6-10 | Budget optimal — configuration économique, adaptée pour fluides non-agressifs (eau, condensat)
Score 11-15 | Standard — bon compromis qualité/prix pour la majorité des applications JESA (acides modérés, eaux industrielles)
Score 16-20 | Premium — configuration haute performance nécessaire pour fluides agressifs (acides forts, slurries)
Score 21+ | Spécial — configuration imposée par les conditions extrêmes (rock slurry, SIL, haute pression). Coût élevé mais pas d'alternative
NOTES IMPORTANTES:
1. Les scores sont des indices relatifs, PAS des prix réels. Score 1 = le moins cher dans sa catégorie, Score 6 = le plus cher.
2. Le TCO réel = CAPEX Score + (OPEX annuel × durée de vie). Les configurations à CAPEX élevé peuvent avoir un OPEX plus bas.
3. Exemple : Config 5 (rock slurry Emerson, score 8) a un CAPEX plus bas que Config 6 (Krohne, score 15), mais le liner soft rubber dure 3-5 ans vs ceramic 10-15 ans.
4. Le modèle AI doit calculer le TCO complet en incluant : nombre de remplacements × coût composant + calibrations × coût calibration + temps d'arrêt × coût horaire.
5. Les durées de vie dans l'onglet OPEX sont des estimations benchmark. Les vraies valeurs nécessitent les données maintenance OCP.
6. Source des scores : JESA Internal DB — Ahmed Saidi 2021, JESA Material Selection DB (PDF), Vendor datasheets (Emerson/E+H/Krohne), Benchmark industrie.

## JESA Project Data
Project | Tag N° | Service (Fluid) | Fluid Type | Line Size (DN) | Line Material | Electrode Material | Liner Material | Grounding Ring Material | Penetrant Ring | O-Ring Material | Flange Material | Flange Coating | Tube Material | Casing Material | Accuracy | Viscosity (cP) | Density (kg/m³) | Normal Flow (m³/h) | Max Flow (m³/h) | Normal Pressure (bar) | Design Pressure (bar) | Normal Temp (°C) | Design Temp (°C) | Max Velocity (m/s) | Remarks | Source
PMP CAP | 204M-FE/FIT-063M | Strong phosphoric acid | Corrosive | 80 | FRP | Tantalum | PFA | 904L SS | N/A | Viton | SS316L | PTFE coating | SS316L | SS316L | ±0.25 to 1% FS | 16 | 1730 | 23.4 | 26 | 8 | 82 | 105 |  | 1 |  | Q37092-00-IC-DAS-00017
PMP CAP | 255-FE/FIT-126M | Sea Water | Corrosive | 500 | FRP | Monel 400 | PFA | Monel 400 | N/A | EPDM | SS316L | None | SS316L | SS316L | ±0.25 to 1% FS | 1 | 1020 | 1200 | 1345 | 3.8 | 7 | 22 | 40 | 1 | Largest meter (DN500) | Q37092-00-IC-DAS-00017
PMP CAP | 204M-FE/FIT-127M | Sea Water | Corrosive | 80 | FRP | Monel 400 | PFA | Monel 400 | N/A | EPDM | SS316L | None | SS316L | SS316L | ±0.25 to 1% FS | 1 | 1020 | 30 | 34 | 0.6 | 7 | 22 | 40 | 2 |  | Q37092-00-IC-DAS-00017
PMP CAP | 213I-FE/FIT-1317 | Weak Phosphoric Acid | Corrosive | 150 | FRP | Tantalum | PFA | 904L SS | N/A | Viton | SS316L | PTFE coating | SS316L | SS316L | ±0.25 to 1% FS | <5 | 1290 | 100 | 112 | 3.8 | 5 | 70 | 105 | 1.8 |  | Q37092-00-IC-DAS-00017
PMP CAP | 213I-FE/FIT-1318 | Weak Phosphoric Acid | Corrosive | 150 | FRP | Tantalum | PFA | 904L SS | N/A | Viton | SS316L | PTFE coating | SS316L | SS316L | ±0.25 to 1% FS | <5 | 1290 | 100 | 112 | 3.8 | 5 | 70 | 105 | 1.8 |  | Q37092-00-IC-DAS-00017
PMP CAP | 213I-FE/FIT-1332 | Weak Phosphoric Acid | Corrosive | 100 | FRP | Tantalum | PFA | 904L SS | N/A | Viton | SS316L | PTFE coating | SS316L | SS316L | ±0.25 to 1% FS | <5 | 1290 | 50 | 58 | 4.8 | 7 | 65 | 105 | 2.1 |  | Q37092-00-IC-DAS-00017
PMP CAP | 213I-FE/FIT-1350 | Weak Phosphoric Acid | Corrosive | 100 | FRP | Tantalum | PFA | 904L SS | N/A | Viton | SS316L | PTFE coating | SS316L | SS316L | ±0.25 to 1% FS | <5 | 1290 | 50 | 58 | 4.8 | 7 | 65 | 105 | 2.1 |  | Q37092-00-IC-DAS-00017
PMP CAP | 213I-FE/FIT-1351 | Weak Phosphoric Acid | Corrosive | 100 | FRP | Tantalum | PFA | 904L SS | N/A | Viton | SS316L | PTFE coating | SS316L | SS316L | ±0.25 to 1% FS | <5 | 1290 | 50 | 58 | 4.8 | 7 | 65 | 105 | 2.1 |  | Q37092-00-IC-DAS-00017
PMP CAP | 213I-FE/FIT-1352 | Weak Phosphoric Acid | Corrosive | 100 | FRP | Tantalum | PFA | 904L SS | N/A | Viton | SS316L | PTFE coating | SS316L | SS316L | ±0.25 to 1% FS | <5 | 1290 | 50 | 58 | 4.8 | 7 | 65 | 105 | 2.1 |  | Q37092-00-IC-DAS-00017
PMP CAP | 213I-FE/FIT-1353 | Weak Phosphoric Acid | Corrosive | 100 | FRP | Tantalum | PFA | 904L SS | N/A | Viton | SS316L | PTFE coating | SS316L | SS316L | ±0.25 to 1% FS | <5 | 1290 | 50 | 58 | 4.8 | 7 | 65 | 105 | 2.1 |  | Q37092-00-IC-DAS-00017
PMP CAP | 255-FE/FIT-1361 | Process Water | Clean | 200 | FRP | Hastelloy C-4 | PFA | Hastelloy C | N/A | Viton | SS316L | None | SS316L | SS316L | ±0.25 to 1% FS | 1 | 1000 | 90 | 100 | 3.8 | 9 | 32 | 60 | 0.8 |  | Q37092-00-IC-DAS-00017
PMP CAP | 213I-FE/FIT-1362 | 98.5% Sulfuric Acid | Corrosive | 50 | ZeCor | Tantalum | PFA | 904L SS | N/A | Viton | SS316L | PTFE coating | SS316L | SS316L | ±0.25 to 1% FS | 14 | 1810 | 2.5 | 3 | 3.8 | 4 | 40 | 60 | 0.4 | Raised Face flange (RF) | Q37092-00-IC-DAS-00017
PMP CAP | 213I-FE/FIT-1340 | 28% Sludge to Reactor | Abrasive | 100 | FRP | Tantalum | PFA | 904L SS | N/A | Viton | SS316L | PTFE coating | SS316L | SS316L | ±0.25 to 1% FS | 200 | 1430 | 35 | 37 |  |  | 73 | 90 | 6 | High viscosity sludge | Q37092-00-IC-DAS-00017
PMP CAP | 213I-FE/FIT-1310 | 28% Sludge to Clarifier | Abrasive | 100 | FRP | Tantalum | PFA | 904L SS | N/A | Viton | SS316L | PTFE coating | SS316L | SS316L | ±0.25 to 1% FS | 200 | 1430 | 35 | 37 |  |  | 73 | 90 | 6 | High viscosity sludge, RF flange | Q37092-00-IC-DAS-00017
Laayoun | N/A | 100% Amine solution | Corrosive | N/A | Metal | Hastelloy C-4 | PTFE | Strap | N/A | EPDM | SS316L | None | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | JESA Flow App 2024
Laayoun | N/A | 5% Amine solution | Corrosive | N/A | Metal | Hastelloy C-4 | PTFE | Strap | N/A | EPDM | SS316L | None | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | JESA Flow App 2024
Laayoun | N/A | 54% Phosphoric Acid | Corrosive | N/A | Metal | Hastelloy C-4 | PTFE | Strap | N/A | Viton | SS316L | PTFE coating | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Hast C-4 instead of Tantalum — project specific | JESA Flow App 2024
Laayoun | N/A | 10% Phosphoric Acid | Corrosive | N/A | Metal | Hastelloy C-4 | PTFE | Strap | N/A | Viton | SS316L | PTFE coating | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Dilute acid — Hast C-4 sufficient | JESA Flow App 2024
Laayoun | N/A | RO water | Clean | N/A | Metal | Hastelloy C-4 | PTFE | Strap | N/A | EPDM | SS316L | None | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | JESA Flow App 2024
Laayoun | N/A | Recovered Process Water | Clean | N/A | Metal | Hastelloy C-4 | PTFE | Strap | N/A | EPDM | SS316L | None | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | JESA Flow App 2024
Laayoun | N/A | 100% Ester solution | Corrosive | N/A | Metal | Hastelloy C-4 | PTFE | Strap | N/A | EPDM | SS316L | None | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | JESA Flow App 2024
Laayoun | N/A | 5% Phosphoric Ester | Corrosive | N/A | Metal | Hastelloy C-4 | PTFE | Strap | N/A | EPDM | SS316L | None | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | JESA Flow App 2024
Laayoun | N/A | Mineral Froth | Abrasive | N/A | FRP | Hastelloy C-4 | PTFE | 2x Hastelloy C-4 | Oui | Viton | SS316L | PTFE coating | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Abrasive + non-conductive pipe | JESA Flow App 2024
Laayoun | N/A | Mineral Slurry | Abrasive | N/A | FRP | Hastelloy C-4 | PTFE | 2x Hastelloy C-4 | Oui | Viton | SS316L | PTFE coating | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Abrasive + non-conductive pipe | JESA Flow App 2024
Laayoun | N/A | Pressurized sea water | Abrasive | N/A | FRP | Hastelloy C-4 | PTFE | 2x Hastelloy C-4 | Oui | Viton | SS316L | PTFE coating | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Abrasive + non-conductive pipe | JESA Flow App 2024
Laayoun | N/A | Sea water | Corrosive | N/A | FRP | Hastelloy C-4 | PTFE | 2x Monel 400 | N/A | EPDM | SS316L | None | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Non-conductive pipe. Hast C-4 electrode — differs from Saidi (Monel 400) and PDF (Titanium) | JESA Flow App 2024
CAP | N/A | 54% Strong H3PO4 | Corrosive | N/A | FRP | Platinum | PFA (SS mesh) | 2x Tantalum | N/A | Viton | SS316L | PTFE coating | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Platinum instead of Tantalum — project specific | JESA Flow App 2024
CAP | N/A | Sea Water | Corrosive | N/A | FRP | Tantalum | PFA (SS mesh) | 2x Monel 400 | N/A | EPDM | SS316L | None | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Tantalum for sea water — unusual choice | JESA Flow App 2024
CAP | N/A | Weak Phosphoric Acid | Corrosive | N/A | N/A | Tantalum | PFA (SS mesh) | Strap | N/A | Viton | SS316L | PTFE coating | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | JESA Flow App 2024
CAP | N/A | Raw Water | Clean | N/A | N/A | Tantalum | PFA (SS mesh) | Strap | N/A | EPDM | SS316L | None | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Tantalum for raw water — over-specified vs Saidi (Alloy C22) | JESA Flow App 2024
PAP Line F | N/A | Sulfuric Acid | Corrosive | N/A | FRP | Tantalum | PFA (SS mesh) | 2x Tantalum | N/A | Viton | SS316L | PTFE coating | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Acid on FRP | JESA Flow App 2024
PAP Line F | N/A | LP Gland Seal Water | Clean | N/A | Metal | Hastelloy C | PFA (SS mesh) | Strap | N/A | EPDM | SS316L | None | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | JESA Flow App 2024
PAP Line F | N/A | MP Gland Seal Water | Clean | N/A | Metal | Hastelloy C | PFA (SS mesh) | Strap | N/A | EPDM | SS316L | None | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | JESA Flow App 2024
PAP Line F | N/A | HP Gland Seal Water | Clean | N/A | Metal | Hastelloy C | PFA (SS mesh) | Strap | N/A | EPDM | SS316L | None | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | JESA Flow App 2024
PAP Line F | N/A | VHP Gland Seal Water | Clean | N/A | Metal | Hastelloy C | PFA (SS mesh) | Strap | N/A | EPDM | SS316L | None | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | JESA Flow App 2024
PAP Line F | N/A | Rock Slurry | Abrasive | N/A | N/A | Conductive rubber | Soft rubber | N/A | Oui | Viton | SS316L | PTFE coating | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Conductive rubber = less noise. Confirms MIA/Daoui lesson | JESA Flow App 2024
PAP Line F | N/A | Weak Phosphoric Acid | Corrosive | N/A | N/A | Tantalum | PFA (SS mesh) | Strap | N/A | Viton | SS316L | PTFE coating | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | JESA Flow App 2024
PAP Line F | N/A | Process Water | Clean | N/A | Metal | Hastelloy C-4 | Soft rubber | Strap | N/A | EPDM | SS316L | None | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Soft rubber for process water — unusual choice | JESA Flow App 2024
PAP Line F | N/A | Raw Water | Clean | N/A | Metal | Alloy C22 | PFA (SS mesh) | Strap | N/A | EPDM | SS316L | None | SS304 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | SS 304 flow tube — cost saving | JESA Flow App 2024
PAP Line F | N/A | Filtered Water | Clean | N/A | Metal | Platinum | PFA (SS mesh) | Strap | N/A | EPDM | SS316L | None | SS304 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Platinum — over-specified vs Saidi (Alloy C22) | JESA Flow App 2024
PAP Line F | N/A | Acidic Process Water | Corrosive | N/A | FRP | Platinum | PFA (SS mesh) | 2x Tantalum | N/A | Viton | SS316L | PTFE coating | SS304 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Platinum instead of Tantalum — project specific | JESA Flow App 2024
PAP Line F | N/A | Reactor Slurry | Abrasive | N/A | Metal | Platinum | PFA (SS mesh) | Strap | N/A | Viton | SS316L | PTFE coating | SS304 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | JESA Flow App 2024
PAP Line F | N/A | Process Condensate | Clean | N/A | Metal | Platinum | PFA (SS mesh) | Strap | N/A | EPDM | SS316L | None | SS304 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Platinum — over-specified vs Saidi (Alloy C22) | JESA Flow App 2024
PAP Line F | N/A | Product Acid | Corrosive | N/A | Metal | Platinum | PFA (SS mesh) | Strap | N/A | Viton | SS316L | PTFE coating | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | JESA Flow App 2024
PAP Line F | N/A | Cake wash water | Clean | N/A | Metal | Platinum | PFA (SS mesh) | Strap | N/A | EPDM | SS316L | None | SS304 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Platinum vs Saidi (Hastelloy C) | JESA Flow App 2024
PAP Line F | N/A | Defoamer | Clean | N/A | Metal | Platinum | PFA (SS mesh) | Strap | N/A | EPDM | SS316L | None | SS304 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | JESA Flow App 2024
PAP Line F | N/A | Sea water | Corrosive | N/A | N/A | Hastelloy C22 | PFA (SS mesh) | Strap | N/A | EPDM | SS316L | None | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Hast C22 for sea water — differs from all other sources | JESA Flow App 2024
SAP | N/A | Caustic Solution | Corrosive | N/A | Metal | Platinum | PTFE | Strap | N/A | Viton | SS304 | None | SS304 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Confirms Saidi (Platinum). Connection Box: CS with Epoxy | JESA Flow App 2024
SAP | N/A | Sulfuric Acid | Corrosive | N/A | FRP | Platinum | PTFE | 2x Tantalum | N/A | Viton | SS305 | PTFE coating | SS304 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Platinum instead of Tantalum — project specific. Connection Box: CS with Epoxy | JESA Flow App 2024
SAP | N/A | Process Water | Clean | N/A | Metal | Platinum | PTFE | Strap | N/A | EPDM | SS316L | None | SS316 | SS304 | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Platinum — over-specified vs Saidi (Hastelloy C-4) | JESA Flow App 2024
DAP | N/A | Weak Phosphoric Acid | Corrosive | N/A | FRP | Tantalum | PFA (SS mesh) | 2x Tantalum | N/A | Viton | SS316L | PTFE coating | SS304L | Die Cast Alu Epoxy | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | JESA Flow App 2024
DAP | N/A | Scrubber Liquor | Corrosive | N/A | Metal | Platinum | PFA (SS mesh) | Strap | N/A | Viton | SS316L | PTFE coating | SS304L | Die Cast Alu Epoxy | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | JESA Flow App 2024
DAP | N/A | Strong Phosphoric Acid | Corrosive | N/A | FRP | Tantalum | PFA (SS mesh) | 2x Tantalum | N/A | Viton | SS316L | PTFE coating | SS304L | Die Cast Alu Epoxy | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | JESA Flow App 2024
DAP | N/A | Sulfuric Acid | Corrosive | N/A | FRP | Tantalum | PFA (SS mesh) | 2x Tantalum | N/A | Viton | SS316L | PTFE coating | SS304L | Die Cast Alu Epoxy | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | JESA Flow App 2024
DAP | N/A | Slurry | Abrasive | N/A | Metal | Platinum | PFA (SS mesh) | Strap | Oui | Viton | SS316L | PTFE coating | SS304L | Die Cast Alu Epoxy | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | JESA Flow App 2024
DAP | N/A | Recovered Liquor | Corrosive | N/A | Metal | Platinum | PFA (SS mesh) | Strap | N/A | Viton | SS316L | PTFE coating | SS304L | Die Cast Alu Epoxy | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | JESA Flow App 2024
DAP | N/A | Raw Water | Clean | N/A | Metal | Alloy C22 | PFA (SS mesh) | Strap | N/A | EPDM | SS316L | None | SS304L | Die Cast Alu Epoxy | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | JESA Flow App 2024
DAP | N/A | Potable Water | Clean | N/A | Metal | Alloy C22 | PFA (SS mesh) | Strap | N/A | EPDM | SS316L | None | SS304L | Die Cast Alu Epoxy | ±0.25 to 1% FS | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | JESA Flow App 2024
TSP Adaptation JFC 2&4 | 507A-FE/FIT-963 | Scrubber Liquor From Primary Scrubber 507AAR02 To Dryer Prescrubber 507AAD06 Inlet — Scrubber Liquor | Corrosive | 80 | 904L SS |  | PFA / PTFE | Grounding ring : (Same as flow tube material) | [6 empty] | +/-0.5% OF FLOW |  |  | 58.5 | 65 |  | Unit | 38 | Unit |  |  | Q36215-JFC2T4-02-IC-DAS-00003
TSP Adaptation JFC 2&4 | 507A-FE/FIT-980 | Scrubber Liquor From Dryer Prescrubber Pumps 507AAP12 & 507AKP12 Outlet To Header — Scrubber Liquor | Corrosive | 50 | 904L SS |  | PFA / PTFE | Grounding ring : (Same as flow tube material) | [6 empty] | +/-0.5% OF FLOW |  |  | 13 | 14 |  | Unit | 39 | Unit |  |  | Q36215-JFC2T4-02-IC-DAS-00003
TSP Adaptation JFC 2&4 | 507A-FE/FIT-960 | Scrubber Liquor From Dryer Prescrubber Pumps 507AAP12 & 507AKP12 TO Dryer Prescrubber 507AAD06 — Scrubber Liquor | Corrosive | 400 | 904L SS |  | PFA / PTFE | Grounding ring : (Same as flow tube material) | [6 empty] | +/-0.5% OF FLOW |  |  | 750 | 1025 |  | Unit | 39 | Unit |  |  | Q36215-JFC2T4-02-IC-DAS-00003
TSP Adaptation JFC 2&4 | 507A-FE/FIT-920 | Strong Phosphoric Acid From Header To Strong Phosphoric Acid Tank 507AAR29 — Strong Phosphoric Acid | Corrosive | 80 | 316/316L SS (DUAL) |  | PFA / PTFE | Grounding ring : (Same as flow tube material) | [6 empty] | +/-0.5% OF FLOW |  |  | 35 | 45 |  | Unit | 39 | Unit |  |  | Q36215-JFC2T4-02-IC-DAS-00003
TSP Adaptation JFC 2&4 | 507A-FE/FIT-922 | Acid Scrubber Liquor Mix From Cone Mixer Feed Pumps 507AAP17 & 507AKP17 Bypass to Strong Phosphoric Acid Tank 507AAR29 — Acid Scrubber Liquor Mix | Corrosive | 80- | 904L SS |  | PFA / PTFE | Grounding ring : (Same as flow tube material) | [6 empty] | +/-0.5% OF FLOW |  |  | 21 | 23.1 |  | Unit | 39 | Unit |  |  | Q36215-JFC2T4-02-IC-DAS-00003
TSP Adaptation JFC 2&4 | 507A-FE/FIT-912 | Acid Scrubber Liquor Mix From Cone Mixer Feed Pump 507AKP17 Outlet To 507AAM05 — Acid Scrubber Liquor Mix | Corrosive | 100 | 904L SS |  | PFA / PTFE | Grounding ring : (Same as flow tube material) | [6 empty] | +/-0.5% OF FLOW |  |  | 54 | 70 |  | Unit | 39 | Unit |  |  | Q36215-JFC2T4-02-IC-DAS-00003
TSP Adaptation JFC 2&4 | 507A-FE/FIT-905 | Acid Scrubber Liquor Mix From Cone Mixer Feed Pump 507AAP17 Outlet To 507AAM05 — Acid Scrubber Liquor Mix | Corrosive | 100 | 904L SS |  | PFA / PTFE | Grounding ring : (Same as flow tube material) | [6 empty] | +/-0.5% OF FLOW |  |  | 54 | 70 |  | Unit | 39 | Unit |  |  | Q36215-JFC2T4-02-IC-DAS-00003
TSP Adaptation JFC 2&4 | 507A-FE/FIT-879 | TSP Slurry From Slurry Pump 507AKP11 To 507AAM02 — Calcium Phosphate Slurry | Clean | 100 | 904L SS |  | PFA / PTFE | Grounding ring : (Same as flow tube material) | [6 empty] | +/-0.5% OF FLOW |  |  | 65 | 80 |  | Unit | 39 | Unit |  |  | Q36215-JFC2T4-02-IC-DAS-00003
TSP Adaptation JFC 2&4 | 507A-FE/FIT-878 | TSP Slurry From Slurry Pump 507AAP11 To 507AAM02 — Calcium Phosphate Slurry | Clean | 100 | 904L SS |  | PFA / PTFE | Grounding ring : (Same as flow tube material) | [6 empty] | +/-0.5% OF FLOW |  |  | 65 | 80 |  | Unit | 39 | Unit |  |  | Q36215-JFC2T4-02-IC-DAS-00003
TSP Adaptation JFC 2&4 | 707A-FE/FIT-963 | Scrubber Liquor From Primary Scrubber 707AAR02 To Dryer Prescrubber 707AAD06 Inlet — Scrubber Liquor | Corrosive | 80 | 904L SS |  | PFA / PTFE | Grounding ring : (Same as flow tube material) | [6 empty] | +/-0.5% OF FLOW |  |  | 58.5 | 65 |  | Unit | 38 | Unit |  |  | Q36215-JFC2T4-02-IC-DAS-00003
TSP Adaptation JFC 2&4 | 707A-FE/FIT-980 | Scrubber Liquor From Dryer Prescrubber Pumps 707AAP12 & 707AKP12 Outlet To Header — Scrubber Liquor | Corrosive | 50 | 904L SS |  | PFA / PTFE | Grounding ring : (Same as flow tube material) | [6 empty] | +/-0.5% OF FLOW |  |  | 13 | 14 |  | Unit | 39 | Unit |  |  | Q36215-JFC2T4-02-IC-DAS-00003
TSP Adaptation JFC 2&4 | 707A-FE/FIT-960 | Scrubber Liquor From Dryer Prescrubber Pumps 707AAP12 & 707AKP12 TO Dryer Prescrubber 707AAD06 — Scrubber Liquor | Corrosive | 400 | 904L SS |  | PFA / PTFE | Grounding ring : (Same as flow tube material) | [6 empty] | +/-0.5% OF FLOW |  |  | 750 | 1025 |  | Unit | 39 | Unit |  |  | Q36215-JFC2T4-02-IC-DAS-00003
TSP Adaptation JFC 2&4 | 707A-FE/FIT-920 | Strong Phosphoric Acid From Header To Strong Phosphoric Acid Tank 707AAR29 — Strong Phosphoric Acid | Corrosive | 80 | 316/316L SS (DUAL) |  | PFA / PTFE | Grounding ring : (Same as flow tube material) | [6 empty] | +/-0.5% OF FLOW |  |  | 35 | 45 |  | Unit | 39 | Unit |  |  | Q36215-JFC2T4-02-IC-DAS-00003
TSP Adaptation JFC 2&4 | 707A-FE/FIT-922 | Acid Scrubber Liquor Mix From Cone Mixer Feed Pumps 707AAP17 & 707AKP17 Bypass to Strong Phosphoric Acid Tank 707AAR29 — Acid Scrubber Liquor Mix | Corrosive | 80- | 904L SS |  | PFA / PTFE | Grounding ring : (Same as flow tube material) | [6 empty] | +/-0.5% OF FLOW |  |  | 21 | 23.1 |  | Unit | 39 | Unit |  |  | Q36215-JFC2T4-02-IC-DAS-00003
TSP Adaptation JFC 2&4 | 707A-FE/FIT-912 | Acid Scrubber Liquor Mix From Cone Mixer Feed Pump 707AKP17 Outlet To 707AAM05 — Acid Scrubber Liquor Mix | Corrosive | 100 | 904L SS |  | PFA / PTFE | Grounding ring : (Same as flow tube material) | [6 empty] | +/-0.5% OF FLOW |  |  | 54 | 70 |  | Unit | 39 | Unit |  |  | Q36215-JFC2T4-02-IC-DAS-00003
TSP Adaptation JFC 2&4 | 707A-FE/FIT-905 | Acid Scrubber Liquor Mix From Cone Mixer Feed Pump 707AAP17 Outlet To 707AAM05 — Acid Scrubber Liquor Mix | Corrosive | 100 | 904L SS |  | PFA / PTFE | Grounding ring : (Same as flow tube material) | [6 empty] | +/-0.5% OF FLOW |  |  | 54 | 70 |  | Unit | 39 | Unit |  |  | Q36215-JFC2T4-02-IC-DAS-00003
TSP Adaptation JFC 2&4 | 707A-FE/FIT-879 | TSP Slurry From Slurry Pump 707AKP11 To 707AAM02 — Calcium Phosphate Slurry | Clean | 100 | 904L SS |  | PFA / PTFE | Grounding ring : (Same as flow tube material) | [6 empty] | +/-0.5% OF FLOW |  |  | 65 | 80 |  | Unit | 39 | Unit |  |  | Q36215-JFC2T4-02-IC-DAS-00003
TSP Adaptation JFC 2&4 | 707A-FE/FIT-878 | TSP Slurry From Slurry Pump 707AAP11 To 707AAM02 — Calcium Phosphate Slurry | Clean | 100 | 904L SS |  | PFA / PTFE | Grounding ring : (Same as flow tube material) | [6 empty] | +/-0.5% OF FLOW |  |  | 65 | 80 |  | Unit | 39 | Unit |  |  | Q36215-JFC2T4-02-IC-DAS-00003
TSP OSBL & Piperack | 310-FE/FIT-401 | 54% PHOSPHORIC ACID FROM FLEXIBILITY LINE JPH 2 FEEDING DAP 107 TO 54% PHOSPHORIC ACID STORAGE TANK 310-A-TK-001 — 54% PHOSPHORIC ACID | Corrosive | 250 | FRP |  | PTFE/PFA | Quantity | [6 empty] | +/-0.2% OF Flow Rate | [5 empty] | Unit |  | Unit |  |  | Q36223-00-00-IC-DAS-00004
TSP OSBL & Piperack | 310-FE/FIT-402 | 54% PHOSPHORIC ACID FROM 54% PHOSPHORIC ACID TRANSFER (DUTY) 310-A-P-001 TO TSP 1 UNIT — 54% PHOSPHORIC ACID | Corrosive | 150 | FRP |  | PTFE/PFA | Quantity | [6 empty] | +/-0.2% OF Flow Rate | [5 empty] | Unit |  | Unit |  |  | Q36223-00-00-IC-DAS-00004
TSP OSBL & Piperack | 310-FE/FIT-403 | 54% PHOSPHORIC ACID FROM 54% PHOSPHORIC ACID TRANSFER (DUTY) 310-B-P-001 TO TSP 2 UNIT — 54% PHOSPHORIC ACID | Corrosive | 150 | FRP |  | PTFE/PFA | Quantity | [6 empty] | +/-0.2% OF Flow Rate | [5 empty] | Unit |  | Unit |  |  | Q36223-00-00-IC-DAS-00004
TSP OSBL & Piperack | 320-FE/FIT-416 | 98% SULFURIC ACID TO SULFURIC ACID STORAGE TANK A — 98% SULFURIC ACID | Corrosive | 150 | Stainless Steel 316/316L (DUAL) |  | PTFE/PFA | Quantity | [6 empty] | +/-0.2% OF Flow Rate | [5 empty] | Unit |  | Unit |  |  | Q36223-00-00-IC-DAS-00004
TSP OSBL & Piperack | 320-FE/FIT-417 | 98% SULFURIC ACID FROM SULFURIC ACID DUTY FEED PUMP 320-A-P-001 TO TSP1 — 98% SULFURIC ACID | Corrosive | 150 | Stainless Steel 316/316L (DUAL) |  | PTFE/PFA | Quantity | [6 empty] | +/-0.2% OF Flow Rate | [5 empty] | Unit |  | Unit |  |  | Q36223-00-00-IC-DAS-00004
TSP OSBL & Piperack | 320-FE/FIT-418 | 98% SULFURIC FROM SULFURIC ACID DUTY FEED PUMP 320-B-P-001 TO TSP2 — 98% SULFURIC ACID | Corrosive | 150 | Stainless Steel 316/316L (DUAL) |  | PTFE/PFA | Quantity | [6 empty] | +/-0.2% OF Flow Rate | [5 empty] | Unit |  | Unit |  |  | Q36223-00-00-IC-DAS-00004
TSP OSBL & Piperack | 330-FE/FIT-454 | RAW WATER FROM RAW WATER TANK 330-A-TK-001 TO PA PUMP 330-A-P-001 — RAW WATER | Clean | 150 | Carbon Steel |  | PTFE/PFA | Quantity | [6 empty] | +/-0.2% OF Flow Rate | [5 empty] | Unit |  | Unit |  |  | Q36223-00-00-IC-DAS-00004
TSP OSBL & Piperack | 530-FE/FIT-460 (Deleted) | POTABLE WATER FROM JPH TP 005 — POTABLE WATER | Clean | 80 | Carbon Steel |  | PTFE/PFA | Quantity | [6 empty] | +/-0.2% OF Flow Rate | [5 empty] | Unit |  | Unit |  |  | Q36223-00-00-IC-DAS-00004
TSP OSBL & Piperack | 530-FE/FIT-461 (Deleted) | POTABLE WATER FROM HEADER TO TSP 1 — POTABLE WATER | Clean | 50 | Carbon Steel |  | PTFE/PFA | Quantity | [6 empty] | +/-0.2% OF Flow Rate | [5 empty] | Unit |  | Unit |  |  | Q36223-00-00-IC-DAS-00004
TSP OSBL & Piperack | 530-FE/FIT-462 (Deleted) | POTABLE WATER FROM HEADER TO TSP 2 — POTABLE WATER | Clean | 50 | Carbon Steel |  | PTFE/PFA | Quantity | [6 empty] | +/-0.2% OF Flow Rate | [5 empty] | Unit |  | Unit |  |  | Q36223-00-00-IC-DAS-00004
TSP OSBL & Piperack | 530-FE/FIT-463 (Deleted) | POTABLE WATER FROM POTABLE WATER TANK TO HEADER — POTABLE WATER | Clean | 80 | Carbon Steel |  | PTFE/PFA | Quantity | [6 empty] | +/-0.2% OF Flow Rate | [5 empty] | Unit |  | Unit |  |  | Q36223-00-00-IC-DAS-00004
Mining 5.0 Beni Amir | 1312-FIT-1301 | Process Water to BT/TBT Scrubber Discharge Pump Box — Process Water | Clean | 450 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.00072 |  | 924.7 | 1063 | 107.7 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1312-FIT-1304 | BT/TBT Scrubber Discharge Pump No.1 / No. 2 Discharge — Process Slurry | Abrasive | 600 | CS ASTM A53 Gr B | Hastelloy C | VTS | Ring & Grounding Strap Hastelloy C | [6 empty] | ± 0.5% of Measured Span | 0.0321 |  | 3675 | 4227 | 417 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1312-FIT-1401 | Process Water to BT/TBT Primary Cyclone Underflow Pump Box — Process Water | Clean | 450 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.00072 |  | 880 | 1012 | 107.4 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1312-FIT-1501 | Process Water to BT/TBT Secondary Cyclone Underflow Pump Box — Process Water | Clean | 300 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.00072 |  | 326 | 375 | 108.2 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1312-FIT-1701 | Process Water to HT/MT Scrubber Discharge Pump Box — Process Water | Clean | 400 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.00072 |  | 725 | 833.7 | 107.3 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1312-FIT-1704 | HT/MT Scrubber Discharge Pump No.1 / No. 2 Discharge — Process Slurry | Abrasive | 600 | CS ASTM A53 Gr B | Hastelloy C | VTS | Ring & Grounding Strap Hastelloy C | [6 empty] | ± 0.5% of Measured Span | 0.0402 |  | 2461 | 2830 | 365 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1312-FIT-1801 | Process Water to HT/MT Primary Cyclone Underflow Pump Box — Process Water | Clean | 350 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.00072 |  | 697 | 802 | 105.2 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1312-FIT-1804 | HT/MT Primary Underflow Pump No.1 / No. 2 Discharge — Process Slurry | Abrasive | 400 | CS ASTM A53 Gr B | Hastelloy C | VTS | Ring & Grounding Strap Hastelloy C | [6 empty] | ± 0.5% of Measured Span | 0.0321 |  | 1202 | 1382 | 414 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1312-FIT-1904 | HT/MT Secondary Cyclone Underflow Pump No.1 / No.2 Discharge — Process Slurry | Abrasive | 350 | CS ASTM A53 Gr B | Hastelloy C | VTS | Ring & Grounding Strap Hastelloy C | [6 empty] | ± 0.5% of Measured Span | 0.0352 |  | 971.9 | 1118 | 449 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1312-FIT-1202 | Process Water to BT/TBT Scrubber Feed Transition Chute — Process Water | Clean | 300 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.00072 |  | 671.2 | 771.9 | 248 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1312-FIT-1206 | BT/TBT Secondary Cyclone Cluster Return Flow — Process Slurry | Abrasive | 500 | CS ASTM A53 Gr B | Hastelloy C | VTS | Ring & Grounding Strap Hastelloy C | [6 empty] | ± 0.5% of Measured Span |  |  | 767.3 | 959.13 |  |  | 35 | 50 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1312-FIT-1602 | Process Water to HT/MT Scrubber — Process Water | Clean | 200 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.00072 |  | 254.5 | 292.6 | 246.7 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1312-FIT-1606 | HT/MT Secondary Cyclone Cluster Return Flow — Process Slurry | Abrasive | 500 | CS ASTM A53 Gr B | Hastelloy C | VTS | Ring & Grounding Strap Hastelloy C | [6 empty] | ± 0.5% of Measured Span |  |  | 578.1 | 722.63 |  |  | 35 | 50 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1313-FIT-2301 | Process Water to BT/TBT Dewatering Cyclone A Feed Pump Box — Process Water | Clean | 350 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.00072 |  | 519 | 597 | 107.2 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1313-FIT-2501 | Process Water to BT/TBT Attritioner A Discharge Pump Box — Process Water | Clean | 350 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.00072 |  | 434 | 499 | 108 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1313-FIT-2504 | BT/TBT Attritioner A Discharge Pump No.1/ No. 2 Discharge — Process Slurry | Abrasive | 400 | CS ASTM A53 Gr B | Hastelloy C | VTS | Ring & Grounding Strap Hastelloy C | [6 empty] | ± 0.5% of Measured Span | 0.0319 |  | 1228 | 1416 | 459 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1313-FIT-2804 | BT/TBT Attritioner B Discharge Pump No.1 /No. 2 Discharge — Process Slurry | Abrasive | 400 | CS ASTM A53 Gr B | Hastelloy C | VTS | Ring & Grounding Strap Hastelloy C | [6 empty] | ± 0.5% of Measured Span | 0.0319 |  | 1231 | 1416 | 515 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1313-FIT-3004 | HT/MT Deslime Circuit Prod Pump No.1 / No. 2 Discharge — Process Slurry | Abrasive | 200 | CS ASTM A53 Gr B | Hastelloy C | VTS | Ring & Grounding Strap Hastelloy C | [6 empty] | ± 0.5% of Measured Span | 0.0406 |  | 285 | 328 | 667.9 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1313-FIT-2206 | Process Water to BT/TBT Combined Deslime Circuit Feed Tank — Process Water | Clean | 350 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.00072 |  | 500 | 575 | 252.7 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1313-FIT-2201A | BT/TBT Deslime Circuit Feed Pump No.1 / No. 2 Discharge — Process Slurry | Abrasive | 600 | CS ASTM A53 Gr B | Hastelloy C | VTS | Ring & Grounding Strap Hastelloy C | [6 empty] | ± 0.5% of Measured Span | 0.0254 |  | 2146 | 2468 | 370 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1313-FIT-2201B | BT/TBT Deslime Circuit Feed Pump No.2 / No. 3 Discharge — Process Slurry | Abrasive | 600 | CS ASTM A53 Gr B | Hastelloy C | VTS | Ring & Grounding Strap Hastelloy C | [6 empty] | ± 0.5% of Measured Span | 0.0254 |  | 2410 | 2771 | 370 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1314-FIT-3301 | Process Water to BT/TBT Ball Mill Discharge Pump — Process Water | Clean | 400 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.00072 |  | 1260 | 1449 | 102.4 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1314-FIT-3307 | BT/TBT Ball Mill Discharge Pump No. 2 / No. 3 Discharge — Process Slurry | Abrasive | 500 | CS ASTM A53 Gr B | Hastelloy C | VTS | Ring & Grounding Strap Hastelloy C | [6 empty] | ± 0.5% of Measured Span | 0.0366 |  | 2017 | 2319 | 533.7 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1314-FIT-4101 | Process Water to HT/MT Ball Mill Discharge Pump — Process Water | Clean | 450 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.00072 |  | 1338 | 1539 | 105.3 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1314-FIT-3707 | BT/TBT Regrind Mill Discharge Pump No.2 / No.3 Discharge — Process Slurry | Abrasive | 300 | CS ASTM A53 Gr B | Hastelloy C | VTS | Ring & Grounding Strap Hastelloy C | [6 empty] | ± 0.5% of Measured Span | 0.0361 |  | 542 | 624 | 514 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1314-FIT-4401 | BT Regrind Blend Pump No.1 / No. 2 Discharge — Process Slurry | Abrasive | 600 | CS ASTM A53 Gr B | Hastelloy C | VTS | Ring & Grounding Strap Hastelloy C | [6 empty] | ± 0.5% of Measured Span | 0.0272 |  | 2574 | 2960 | 446 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1314-FIT-3401 | Process Water to BT/TBT Ball Mill A Feed Transition Chute — Process Water | Clean | 150 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.00072 |  | 150 | 172.5 | 242.9 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1314-FIT-4201 | Process Water to HT/MT Ball Mill A Feed Transition Chute — Process Water | Clean | 150 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.00072 |  | 150 | 173 | 242.9 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1314-FIT-4301 | Process Water to HT/MT Ball Mill B Feed Transition Chute — Process Water | Clean | 150 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.00072 |  | 150 | 172.5 | 242.9 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1314-FIT-3609 | Process Water to BT/TBT Product Combination Tank — Process Water | Clean | 150 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.00072 |  | 150 | 173 | 242.9 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1314-FIT-3603 | BT/TBT Prod Thickener Feed Pump No. 2 / No. 3 Discharge — Process Slurry | Abrasive | 350 | CS ASTM A53 Gr B | Hastelloy C | VTS | Ring & Grounding Strap Hastelloy C | [6 empty] | ± 0.5% of Measured Span | 0.0333 |  | 1105 | 1270 | 404 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1314-FIT-3601 | BT/TBT Regrind Circuit Feed Pump No.1 / No. 2 Discharge — Process Slurry | Abrasive | 200 | CS ASTM A53 Gr B | Hastelloy C | VTS | Ring & Grounding Strap Hastelloy C | [6 empty] | ± 0.5% of Measured Span | 0.0334 |  | 279.1 | 321 | 271 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1315-FIT-4603 | Raw Water to BT/TBT Flotation Circuit A Feed Dilution Box — Raw Water | Clean | 200 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.00072 |  | 215.4 | 247.7 | 239.2 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1315-FIT-5203 | Raw Water to BT/TBT Flotation Circuit B Feed Dilution Box — Raw Water | Clean | 200 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.00072 |  | 215.4 | 247.7 | 239.2 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1315-FIT-5101 | BT/TBT Flotation Circuit A Cyclone Feed Pump No. 1/ No. 2 Discharge — Process Slurry | Abrasive | 500 | CS ASTM A53 Gr B | Hastelloy C | VTS | Ring & Grounding Strap Hastelloy C | [6 empty] | ± 0.5% of Measured Span | 0.0279 |  | 2102 | 2417 | 457.3 |  | 35 | 50 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1315-FIT-4601 | Raw Water to BT/TBT Flotation Circuit A Conditioner Tank 1 — Raw Water | Clean | 250 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.00072 |  | 344 | 396 | 240.6 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1316-FIT-6503 | Tails Thickener Inline Flocculant Mixer Outlet — Flocculant | Clean | 150 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.00072 |  | 72.5 | 83.4 | 137 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1316-FIT-9401 | Tails Thickener A Underflow Pump No.1 / No.2 Discharge — Process Slurry | Abrasive | 350 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.0468 |  | 826 | 950 | 550 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1316-FIT-5903 | BT/TBT Thickener Flocculant Mixer Outlet — Flocculant | Clean | 150 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span |  |  | 76.4 | 100 | 138 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1316-FIT-6001 | BT/TBT Thickener Product Flow post pumps — Process Slurry | Abrasive | 350 | CS ASTM A53 Gr B | Hastelloy C | VTS | Ring & Grounding Strap Hastelloy C | [6 empty] | ± 0.5% of Measured Span |  |  | 738.3 | 922.88 |  |  | 35 | 50 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1316-FIT-6201 | HT/MT Thickener Underflow Pump No. 1/ No.2 Discharge — Process Slurry | Abrasive | 350 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.0363 |  | 792 | 910 | 415 |  | 35 | 50 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1318-FIT-7502 | Raw Water Pump Discharge — Raw Water | Clean | 600 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.00072 |  | 2000 | 2465 | 160 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1316-FIT-9801 | Combined Tails Transfer Pump No. A2/ C2 Discharge — Process Slurry | Abrasive | 750 | CS ASTM A53 Gr B | Hastelloy C | VTS | Ring & Grounding Strap Hastelloy C | [6 empty] | ± 0.5% of Measured Span | 0.06 |  | 1677 | 2013 | 1200 |  | 35 | 50 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1318-FIT-7501 | Raw Water Pump No.1/ No.2/ No.3/ No. 4 Discharge — Raw Water | Clean | 750 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.00072 |  | 1185 | 1362 | 324 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1318-FIT-7202 | Raw Water to Process Water Pond — Raw Water | Clean | 400 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.00072 |  | 700 | 805 | 146.5 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1318-FIT-7306B | Process Water Pump No. 4 to No. 6 Discharge — Raw Water | Clean | 1050 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.00072 |  | 6304 | 7250 | 187.4 |  | 20 | 30 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1317-FIT-6804 | Amine Mixing Tank Raw Water Flow Measurement — Raw Water | Clean | 200 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span | 0.00072 |  | 245.2 | 306.5 |  |  | 35 | 50 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1317-FIT-6803 | Amine Mixing Tank Dosing Flow Measurement — Amine | Corrosive | 150 | SS ASTM A312-TP316L | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span |  |  | 29.6 | 37 |  |  | 40 | 50 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1317-FIT-6904 | Phosphoric Acid Mixing Tank Raw Water Flow Measurement — Raw Water | Clean | 200 | CS ASTM A53 Gr B | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span |  |  | 183.9 | 229.88 |  |  | 35 | 50 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1317-FIT-6903 | Phosphoric Acid Mixing Tank Dosing Flow Measurement — Phosphoric Ester | Corrosive | 150 | SS ASTM A312-TP316L | SS 316 | VTS | Grounding Strap SS316 | [6 empty] | ± 0.5% of Measured Span |  |  | 22.2 | 27.75 |  |  | 40 | 50 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1315-FIT-4755 | Flotation cells FC001-Product Flow — Slurry | Abrasive |  |  | Hastelloy C | VTS | Ring & Grounding Strap Hastelloy C | [6 empty] | ± 0.5% of Measured Span |  |  | 2028.8 | 2536 |  |  | 35 | 50 |  |  | Q12203-01-IC-DAS-00005
Mining 5.0 Beni Amir | 1315-FIT-5355 | Flotation cells FC009-Product Flow — Slurry | Abrasive |  |  | Hastelloy C | VTS | Ring & Grounding Strap Hastelloy C | [6 empty] | ± 0.5% of Measured Span |  |  | 2757 | 3446.25 |  |  | 35 | 50 |  |  | Q12203-01-IC-DAS-00005
CAP SAFI Existing Site | E10-FIT-401 | 54% Phosphoric Acid from Storage Tanks MP1 to SPH — 54% Phosphoric Acid | Corrosive | 400 | 904L SS | Platinum | PFA | Grounding Straps SS316L | [6 empty] | ± 0.2% of Flowrate | 80 | 1810 | 450 | 495 | 17 | 24.5 | 35 | 80 | 1.81 | Tube: 316L SS, Flange: 316L SS, 300# RF | Q80130-CSP-E10-IC-DAS-0002
CAP SAFI Existing Site | E10-FIT-404 | 54% Phosphoric Acid from Storage Tanks MP1 to SPH — 54% Phosphoric Acid | Corrosive | 400 | 904L SS | Platinum | PFA | Grounding Straps SS316L | [6 empty] | ± 0.2% of Flowrate | 80 | 1810 | 450 | 495 | 17 | 24.5 | 35 | 80 | 1.81 | Tube: 316L SS, Flange: 316L SS, 300# RF | Q80130-CSP-E10-IC-DAS-0002
Jorf/Safi Desalinated Water Pipelines | 187-FIT-911 | Flow measurement in water pipeline — Potable Water | Clean | 1600 | FRP (Fiberglass) | SS 316L | Hard Rubber / Ebonite | Rings or electrodes SS 316L | [6 empty] | 0.5% of full scale |  |  |  | 12408 |  | 1.2 | 20.5 | 22 |  | DN 1600, 150# FF AWWA C207 Class D, FRP pipe, 220V AC | QW212201-04-IC-DAS-00001-002
UPDATED SUMMARY
Original entries: 58 ¦ New entries: 84 ¦ Grand total: 142
CAP SAFI Existing Site | 2 magflows
Jorf/Safi Desalinated Water Pipelines | 1 magflows
Mining 5.0 Beni Amir | 52 magflows
TSP Adaptation JFC 2&4 | 18 magflows
TSP OSBL & Piperack | 11 magflows